## device_event_limit25
report_number | manufacturer_contact_state | device_date_of_manufacturer | event_type | source_type- | source_type001 | source_type002 | manufacturer_g1_address_1 | manufacturer_contact_pcountry | report_to_manufacturer | previous_use_code | manufacturer_postal_code | manufacturer_city | reporter_occupation_code | manufacturer_contact_extension | date_facility_aware | manufacturer_state | event_location | date_of_event | manufacturer_g1_country | removal_correction_number | manufacturer_g1_city | mdr_text__text | mdr_text__mdr_text_key | mdr_text__text_type_code | mdr_text__patient_sequence_number | distributor_zip_code_ext | date_received | manufacturer_zip_code | manufacturer_contact_address_2 | date_report | manufacturer_contact_address_1 | manufacturer_contact_plocal | manufacturer_name | type_of_report- | type_of_report001 | type_of_report002 | reprocessed_and_reused_flag | manufacturer_link_flag | manufacturer_contact_zip_code | manufacturer_g1_zip_code_ext | manufacturer_contact_country | distributor_zip_code | report_source_code | distributor_name | date_manufacturer_received | remedial_action- | remedial_action | manufacturer_contact_postal_code | manufacturer_contact_city | product_problem_flag | event_key | manufacturer_g1_zip_code | manufacturer_g1_state | patient__date_received | patient__patient_sequence_number | patient__sequence_number_treatment | patient__sequence_number_outcome001 | patient__sequence_number_outcome002 | patient__sequence_number_outcome- | manufacturer_contact_phone_number | manufacturer_zip_code_ext | manufacturer_contact_pcity | manufacturer_g1_postal_code | manufacturer_contact_area_code | adverse_event_flag | single_use_flag | distributor_city | manufacturer_country | manufacturer_contact_f_name | distributor_state | number_devices_in_event | manufacturer_contact_exchange | distributor_address_2 | manufacturer_contact_t_name | distributor_address_1 | manufacturer_g1_address_2 | health_professional | initial_report_to_fda | number_patients_in_event | mdr_report_key | report_to_fda | manufacturer_address_2 | report_date | manufacturer_address_1 | manufacturer_g1_name | manufacturer_contact_zip_ext | manufacturer_contact_l_name | device__date_removed_flag | device__other_id_number | device__manufacturer_d_zip_code_ext | device__baseline_510_k__number | device__implant_flag | device__device_operator | device__brand_name | device__lot_number | device__expiration_date_of_device | device__manufacturer_d_state | device__generic_name | device__manufacturer_d_name | device__device_availability | device__manufacturer_d_city | device__device_sequence_number | device__device_age_text | device__catalog_number | device__baseline_510_k__exempt_flag | device__date_received | device__manufacturer_d_address_2 | device__manufacturer_d_address_1 | device__manufacturer_d_postal_code | device__baseline_510_k__flag | device__model_number | device__manufacturer_d_zip_code | device__device_report_product_code | device__date_returned_to_manufacturer | device__manufacturer_d_country | device__device_event_key | device__device_evaluated_by_manufacturer
3009450871-2014-10024 | PA | 2006-03-25 | Malfunction |  | Health Professional | Company representation | EIMATTSTRASSE 3 |  | N | * | CH4436 | OBERDORF | SERVICE AND TESTING PERSONNEL |  |  |  |  | 20140225 | SZ |  | OBERDORF | THE ACTUAL DEVICE HAS BEEN RETURNED AND IS CURRENTLY PENDING EVALUATION. ONCE RELIABILITY ENGINEERING EVALUATES THE DEVICE, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SENT ACCORDINGLY. | 24541745 | Additional Manufacturer Narrative | 1 |  | 20140327 |  |  | 20140225 | 1302 WRIGHTS LANE EAST |  | SYNTHES OBERDORF |  | Initial submission | Followup | N | Y | 19380 |  | US |  | Manufacturer report |  | 20140724 |  |  |  | WEST CHESTER | Y | 3602907 |  |  | 20140327 | 1 |  |  |  |  | 5000 |  |  | CH4436 | 610 | N | N |  | SZ | LINDA |  | 1 | 719 |  |  |  |  | N | No | 1 | 3706369 |  |  |  | EIMATTSTRASSE 3 | SYNTHES OBERDORF |  | PLEWS | B |  |  |  | N | HEALTH PROFESSIONAL | MAINTENANCE STATION F/PEN DRIVE | 5205351 |  |  | INSTR, SURGICAL, ORTHOPEDIC, AC- POWERED MOTOR/ACCESS & ATTACH | SYNTHES OBERDORF | Device was returned to manufacturer | OBERDORF | 1 | DA | 05.001.099 |  | 20140327 |  | EIMATTSTRASSE 3 | CH4436 |  |  |  | HWE | 20140225 | SZ | 3727602 | Y
[22 empty] | IT WAS REPORTED THAT DURING SERVICE AND REPAIR PRE-TESTING, IT WAS DISCOVERED THAT THE ELECTRIC PEN DRIVE DEVICE ¿DID NOT STOP AFTER ONE MINUTE¿. THE EVENT WAS NOT RELATED TO SURGERY. THERE WERE NO INJURIES OR MEDICAL INTERVENTION ASSOCIATED WITH THIS EVENT. ALL AVAILABLE INFORMATION HAS BEEN DISCLOSED. IF ADDITIONAL INFORMATION SHOULD BECOME AVAILABLE, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED ACCORDINGLY. | 24541740 | Description of Event or Problem | 1
[22 empty] | ADDITIONAL NARRATIVE: CORRECTION: THE MANUFACTURER LOCATION WAS DOCUMENTED AS WALDENBURG IN THE INITIAL REPORT. THE LOCATION HAS BEEN UPDATED TO (B)(4). CONTACT OFFICE NAME/ADDRESS HAVE BEEN UPDATED ACCORDINGLY TO REFLECT THE CORRECTED MANUFACTURING FACILITY. DEVICE EVALUATION: THE DEVICE WAS RETURNED FOR SERVICE HOWEVER DID NOT MEET MANUFACTURING SPECIFICATIONS DURING PRE-REPAIR ASSESSMENT. RELIABILITY ENGINEERING EVALUATED THE DEVICE AND THE REPORTED CONDITION WAS CONFIRMED. THE ASSIGNABLE ROOT CAUSE WAS DETERMINED TO BE MISUSE, ABUSE AND/OR ERROR. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL MEDWATCH, A FOLLOW-UP MEDWATCH WILL BE FILED AS APPROPRIATE. | 31673099 | Additional Manufacturer Narrative | 1
1045834-2013-15672 | FL | 2008-11-01 | Malfunction | User facility |  |  |  |  | N | U |  |  | OTHER |  |  |  |  | 20110615 |  |  |  | THE DEVICE HAS BEEN RECEIVED BY DEPUY SYNTHES POWER TOOLS. THE INVESTIGATION IS STILL IN PROCESS. WHEN THE INVESTIGATION HAS BEEN COMPLETED OR ADDITIONAL INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SENT. (B)(4). | 25874353 | Additional Manufacturer Narrative | 1 |  | 20131211 |  |  | 20110615 | 4500 RIVERSIDE DR. |  | DEPUY SYNTHES POWER TOOLS | Initial submission |  |  | N | Y | 33410 |  | US |  | Manufacturer report |  | 20110615 |  |  |  | PALM BEACH GARDENS | Y | 3871796 |  |  | 20131211 | 1 | [9 empty] | N | N |  |  | RODNEY |  | 1 | [5 empty] | N | Unknown | 1 | 3975201 |  |  |  | 4500 RIVERSIDE DRIVE |  |  | CRAIN | B |  |  |  | N | HEALTH PROFESSIONAL | EMAX 2 PLUS MOTOR |  |  |  |  | DEPUY SYNTHES POWER TOOLS | Yes |  | 1 | DA | EMAX2PLUS_LL |  | 20131211 |  | 4500 RIVERSIDE DRIVE |  |  |  |  | HBC | 20110621 |  | 4024786 | N
[22 empty] | REPORT RECEIVED FROM THE USA STATING THAT THE DEVICE WOULD NOT RUN AND HAD STOPPED WORKING DURING THE CASE. IT IS KNOWN THE DEVICE WAS USED DURING SURGERY. IT IS KNOWN NO INJURY OR MEDICAL INTERVENTION OCCURRED. THERE WAS NO ADDITIONAL INFORMATION PROVIDED. | 25874363 | Description of Event or Problem | 1
1031452-2014-16923 | OH |  | Malfunction | Other |  |  | 2101 EAST LAKE MARY BLVD |  | N | R |  | SANFORD | MEDICAL EQUIPMENT COMPANY TECHNICIAN/REPRESENTATIVE |  |  | FL |  |  | US |  | SANFORD | INDEPENDENT REPAIR CENTER CUSTOMER ALLEGED, UNIT IS ALARMING OR RED LIGHT AND THE KEY FAILURE IS THE SIEVE BEDS ARE DUSTED. | 27746198 | Description of Event or Problem | 1 |  | 20141030 | 32773 |  | 20141009 | ONE INVACARE WAY |  | INVACARE FLORIDA OPERATIONS | Initial submission |  |  | N | Y | 44035 |  | US |  | Manufacturer report |  | 20141009 |  |  |  | ELYRIA | Y | 4112468 | 32773 | FL | 20141030 | 1 |  |  |  | 1, Other | 6900 |  |  |  | 800 | N | N |  | US | KAREN |  | 1 | 333 |  |  |  |  | N | No | 1 | 4215873 |  |  |  | 2101 EAST LAKE MARY BLVD | INVACARE FLORIDA OPERATIONS |  | LOUGHREN | B |  |  |  | N | LAY USER/PATIENT | PORTABLE OXYGEN GENERATOR (CONCENTRATOR, HOMEFILL) |  |  | FL | 868.544 | INVACARE FLORIDA OPERATIONS | No | SANFORD | 1 |  |  |  | 20141030 |  | 2101 EAST LAKE MARY BLVD |  |  | IRC5PO2V | 32773 | CAW |  | US | 4245517 | R
2210968-2013-23465 | OH |  | Injury | Other |  |  | 982 ROAD 183 KM 8.3 |  | N | I |  | SOMERVILLE | OTHER |  |  | NJ |  |  |  | NA | SAN LORENZO | (B)(4). CONCLUSION: NO CONCLUSION CAN BE DRAWN AT THIS TIME. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A SUPPLEMENTAL 3500A FORM WILL BE SUBMITTED ACCORDINGLY. | 20617972 | Additional Manufacturer Narrative | 1 |  | 20130918 | 8876 |  | 20130911 | 4545 CREEK RD |  | ETHICON INC. |  | Initial submission | Followup | N | Y | 45242 |  | US |  | Manufacturer report |  | 20150319 |  |  |  | CINCINNATI | N | 3252128 |  | PR | 20130918 | 1 |  |  |  | 1, Required Intervention | 3299 | 151 |  | 754 | 513 | Y | Y |  | US | KATHY |  | 1 | 337 |  |  |  |  | N | Unknown | 1 | 3355553 | N |  |  | P.O. BOX 151, ROUTE 22 WEST | ETHICON INC. - SAN LORENZO PR | 2803 | RICE | B |  | 151 |  | Y | HEALTH PROFESSIONAL | INTERCEED* (TC7) ABSORBABLE ADHESION BARRIER | ZBB108 |  | NJ | BARRIER, ABSORBABLE, ADHESION | ETHICON INC. | Yes | SOMERVILLE | 1 | DA | UNK |  | 20130918 |  | P.O. BOX 151, ROUTE 22 WEST |  |  | UNK | 8876 | MCN |  | US | 3383948 | N
[22 empty] | IT WAS REPORTED THAT THE PATIENT UNDERWENT A GYNECOLOGICAL PROCEDURE ON (B)(6) 2007 AND A MESH WAS IMPLANTED. IT WAS REPORTED THAT SHE EXPERIENCED PAIN, EROSION OF HER INTERNAL BODILY TISSUE AND OTHER INJURIES FOLLOWING THE PROCEDURE. IT WAS REPORTED THAT THE PATIENT HAS UNDERGONE MULTIPLE SURGERIES AND REVISIONARY PROCEDURES. NO ADDITIONAL INFORMATION WAS PROVIDED. | 20617970 | Description of Event or Problem | 1
[22 empty] | IT WAS REPORTED BY AN ATTORNEY THAT THE PATIENT UNDERWENT A GYNECOLOGICAL PROCEDURE ON (B)(6) 2015 AND A MESH WAS IMPLANTED CONCURRENTLY WITH LAPAROSCOPY WITH LYSIS OF ADHESIONS, AND EXPLORATORY LAPAROSCOPY WITH LYSIS OF ADHESIONS AND LEFT SALPINGO-OOPHORECTOMY. NO ADDITIONAL INFORMATION WAS PROVIDED. | 30552711 | Additional Manufacturer Narrative | 1
1818910-2013-02933 | IN | 2007-02-10 | Injury | Other |  |  | ST. ANTHONYS RD |  | N | I |  | LEEDS | ATTORNEY | [5 empty] | UK | Z1749/1816-2011 | LEEDS | LITIGATION ALLEGES PATIENT HAD PAIN, WEAKNESS AND DIFFICULTY WITH SIMPLE DAILY LIVING ACTIVITIES AFTER ASR HIP IMPLANT. | 17710000 | Description of Event or Problem | 1 |  | 20130301 |  |  | 20130226 | 700 ORTHOPAEDIC DRIVE |  | DEPUY INTERNATIONAL |  | Followup | Initial submission | N | Y | 46581 |  | US |  | Manufacturer report |  | 20130326 | Recall |  |  | WARSAW | N | 2881571 |  |  | 20130401 | 1 |  |  |  | 1, Other | [5 empty] | Y | Y |  | UK | STACEY |  | 1 | [5 empty] | N | No | 1 | 2985057 |  |  |  | ST. ANTHONYS RD | DEPUY INTERNATIONAL | 988 | TRICK | B |  |  |  | Y | HEALTH PROFESSIONAL | ASR UNI FEMORAL IMPL SIZE 47 | 2319442 |  |  | TOTAL HIP REPLACEMENT | DEPUY INTERNATIONAL | No | LEEDS | 1 | DA | 999890247 |  | 20130301 |  | ST. ANTHONYS RD |  |  |  |  | KWA |  | UK | 3010652 | R
[22 empty] | THE ASR PLATFORM WAS VOLUNTARILY RECALLED FROM THE MARKET IN AUGUST 2010, AND THE ASR PRODUCT CODES ARE NOW CONSIDERED INACTIVE. FURTHER INVESTIGATION OF THIS INDIVIDUAL INCIDENT WILL NOT BE UNDERTAKEN, AS THERE IS AN ONGOING INVESTIGATION REGARDING THE ROOT CAUSE(S) AND/OR CORRECTIVE ACTIONS. (B)(4).DEPUY CONSIDERS THE INVESTIGATION CLOSED AT THIS TIME. SHOULD THE PRODUCT AND/OR ADDITIONAL INFORMATION BE RECEIVED, THE INVESTIGATION WILL BE RE-OPENED. | 17710001 | Additional Manufacturer Narrative | 1
[22 empty] | LITIGATION ALLEGES PATIENT HAD PAIN, WEAKNESS AND DIFFICULTY WITH SIMPLE DAILY LIVING ACTIVITIES AFTER ASR HIP IMPLANT. UPDATE - (B)(4) 2013 - PLAINTIFF'S PRELIMINARY DISCLOSURE FORM WAS RECEIVED, WHICH IDENTIFIED PART/LOT INFORMATION. THE COMPLAINT AND ASSOCIATED MDRS WERE UPDATED. THERE WAS NO NEW INFORMATION THAT WOULD CHANGE THE OUTCOME OF THE INVESTIGATION. | 18087594 | Description of Event or Problem | 1
[22 empty] | CORRECTED: EVENT/PROBLEM DESCRIPTION, BRAND NAME, DEVICE PRODUCT CODE, CATALOG #/LOT #, DATE RECEIVED BY MANUFACTURER, PMA/510(K) #, MANUFACTURE DATE. DEPUY STILL CONSIDERS THIS INVESTIGATION CLOSED. | 18087597 | Additional Manufacturer Narrative | 1
1031452-2014-11164 | OH |  | Malfunction | Other |  |  | 2101 EAST LAKE MARY BLVD |  | N | R |  | SANFORD | MEDICAL EQUIPMENT COMPANY TECHNICIAN/REPRESENTATIVE |  |  | FL |  |  | US |  | SANFORD | PER INDEPENDENT REPAIR CENTER: WILL NOT START SHORT CIRCUIT. | 26938651 | Description of Event or Problem | 1 |  | 20140926 | 32773 |  | 20140904 | ONE INVACARE WAY |  | INVACARE FLORIDA OPERATIONS | Initial submission |  |  | N | Y | 44035 |  | US |  | Manufacturer report |  | 20140904 |  |  |  | ELYRIA | Y | 4018568 | 32773 | FL | 20140926 | 1 |  |  |  | 1, Other | 6900 |  |  |  | 800 | N | N |  | US | KAREN |  | 1 | 333 |  |  |  |  | N | No | 1 | 4121953 |  |  |  | 2101 EAST LAKE MARY BLVD | INVACARE FLORIDA OPERATIONS |  | LOUGHREN | B |  |  |  | N | LAY USER/PATIENT | PORTABLE OXYGEN GENERATOR (CONCENTRATOR, HOMEFILL) |  |  | FL | 868.544 | INVACARE FLORIDA OPERATIONS | No | SANFORD | 1 |  |  |  | 20140926 |  | 2101 EAST LAKE MARY BLVD |  |  | IRC10LXO2 | 32773 | CAW |  | US | 4149688 | R
3004209178-2013-15741 | MN | 2007-04-20 | Malfunction | Health Professional |  |  | 7000 CENTRAL AVENUE NE RCW215 |  | N | I |  | JUNCOS | PHYSICIAN |  |  | PR |  | 20130814 | US |  | MINNEAPOLIS | IT WAS REPORTED THAT THE PATIENT¿S PUMP WAS FILLED WITH INCORRECT MEDICATION. THE PHARMACY SENT THE HEALTHCARE PROVIDER¿S (HCP) OFFICE MEDICATION THAT WAS NOT PRESERVATIVE FREE AND THAT CONTAINED DEXTROSE FOR A PUMP REFILL. THE HCP CALLED THE PATIENT INTO THE OFFICE ON 2013 (B)(6) AND CHANGED THE MEDICATION OUT AFTER PERFORMING A RESERVOIR RINSE. NO PATIENT SYMPTOMS WERE REPORTED. THE DEVICE SYSTEM DELIVERED BACLOFEN, DILAUDID AND BUPIVACAINE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT WAS NOT AVAILABLE AT THE TIME OF THIS REPORT. | 22237229 | Description of Event or Problem | 1 |  | 20130905 | 777 |  | 20130814 | 7000 CENTRAL AVENUE NE RCW215 |  | MDT PUERTO RICO OPERATIONS CO | Initial submission |  |  | N | Y | 55432 |  | US |  | Manufacturer report |  | 20130814 |  |  |  | MINNEAPOLIS | Y | 3223332 | 55432 | MN | 20130905 | 1 | .; |  |  |  | 9619 |  |  |  | 763 | N | Y |  | US | CHAD |  | 1 | 514 |  |  |  |  | Y | Unknown | 1 | 3328033 |  |  |  | RD 31 KM 24 HM 4 | MEDTRONIC NEUROMODULATION |  | HEDLUND | B |  |  |  | Y | HEALTH PROFESSIONAL | SYNCHROMED II |  | 20081014 | PR | PUMP, INFUSION, IMPLANTED, PROGRAMMABLE | MDT PUERTO RICO OPERATIONS CO | No | JUNCOS | 1 |  | 863720 |  | 20130905 |  | RD 31 KM 24 HM 4 |  |  | 863720 | 777 | LKK |  | US | 3357507 | *
[22 empty] | PRODUCT ID 8709, SERIAL# (B)(4), IMPLANTED: 2007 (B)(6); PRODUCT TYPE CATHETER. (B)(4). | 22237264 | Additional Manufacturer Narrative | 1
1314492-2014-44518 | NY | 2011-05-18 | Malfunction |  | Health Professional | Company representation |  |  | N | R |  | MEDINA | BIOMEDICAL ENGINEER |  |  | NY |  | 20140601 |  |  |  | DURING THE EVAL OF A RETURNED SPECTRUM INFUSION PUMP, 13 OCCURRENCES OF 'UPSTREAM OCCLUSION' ALARMS WERE IDENTIFIED IN THE EVENT HISTORY LOG. ANY PT INVOLVEMENT, INJURY OR MEDICAL INTERVENTION IS UNK SINCE THESE ITEMS WERE FOUND DURING EVAL OF THE DEVICE. | 29841563 | Description of Event or Problem | 1 |  | 20140717 |  |  | 20140702 | 711 PARK AVE. |  | BAXTER HEALTHCARE CORP | Initial submission |  |  | N | Y | 14103 |  | US |  | Manufacturer report |  | 20140702 | Repair |  |  | MEDINA | Y | 4347067 |  |  | 20140717 | 1 |  |  |  |  | 3454 |  |  |  | 800 | N | N |  | US | KELLY |  | 1 | 356 |  |  |  |  | Y | Unknown | 1 | 4476193 | [6 empty] | MORSE, SR MGR, QUAL | B |  |  |  | N | HEALTH PROFESSIONAL | SPECTRUM INFUSION PUMP |  |  | NY | INFUSION PUMP | BAXTER HEALTHCARE CORP | Yes | MEDINA | 1 |  | 35700 |  | 20140717 |  |  |  |  | 35700 |  | FRN | 20140627 | US | 4552970 | Y
[22 empty] | (B)(4). THE DEVICE WAS RETURNED AND EVALUATED BY BAXTER. THIS COMPLAINT WAS OPENED DURING THE INVESTIGATION OF COMPLAINT (B)(4). THE UPSTREAM OCCLUSION ALARMS WERE CONFIRMED THROUGH AN EVENT HISTORY LOG REVIEW. THE CAUSE WAS UNDETERMINED. IF ANY ADD'L INFO IS REC'D A FOLLOW UP WILL BE SENT. THE ULTRASONIC SENSOR AND ULTRASONIC PUSHER WERE REPLACED. | 29841578 | Additional Manufacturer Narrative | 1
2649622-2014-13557 | MN | 2013-03-08 | Malfunction |  | Health Professional | Company representation | 8200 CORAL SEA ST NE |  | N | I |  | VILLALBA | PHYSICIAN |  |  | PR |  | 20140907 | US |  | MOUNDS VIEW | THIS EVENT OCCURRED OUTSIDE THE U.S. WHERE A SIMILAR MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. (B)(4). | 28623009 | Additional Manufacturer Narrative | 1 |  | 20141210 | 766 |  | 20140908 | 8200 CORAL SEA ST NE |  | MPRI | Initial submission |  |  | N | Y | 55112 |  | US |  | Manufacturer report |  | 20140908 |  |  |  | MOUNDS VIEW | Y | 4208691 | 55112 | MN | 20141210 | 1 |  |  |  |  | 2036 |  |  |  | 763 | N | Y |  | US | ANNE |  | 1 | 505 |  |  |  |  | Y | Unknown | 1 | 4312097 |  |  |  | ROAD 149 KM 56.3 | MEDTRONIC CARDIAC RHYTHM HEART FAILURE |  | SCHILLING | V |  |  |  | N | HEALTH PROFESSIONAL | CAPSUREFIX MRI SURESCAN |  | 20150226 | PR | DRUG ELUTING PERMANENT RIGHT VENTRICULAR (RV) OR RIGHT ATRIAL (RA) PACEMAKER ELE | MPRI | No | VILLALBA | 1 |  | 5086MRI45 |  | 20141210 |  | ROAD 149 KM 56.3 |  |  | 5086MRI45 | 766 | NVN |  | US | 4341775 | *
[22 empty] | IT WAS REPORTED THAT DURING AN IMPLANT ATTEMPT, DURING POSITIONING OF THE LEAD THE HELIX COULD NOT BE EXTRACTED EASILY. THE LEAD WAS REMOVED FROM THE PATIENT TO TEST EXTRACTION/RETRACTION OF THE HELIX. IT WAS NOT EXTRACTED EVEN AFTER 20 ROTATIONS AND THEREFORE CONSIDERED ABNORMAL. ANOTHER LEAD WAS IMPLANTED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. | 28623000 | Description of Event or Problem | 1
2210968-2013-17244 | OH |  | Injury | Other |  |  | PUITS-GODET 20 2000 NEUCHATEL |  | N | I |  | SOMERVILLE | OTHER |  |  | NJ |  |  | SZ | NA | NEUCHATEL | (B)(4). CONCLUSION: NO CONCLUSION CAN BE DRAWN AT THIS TIME. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A SUPPLEMENTAL 3500A FORM WILL BE SUBMITTED ACCORDINGLY. | 20023784 | Additional Manufacturer Narrative | 1 |  | 20130809 | 8876 |  | 20130718 | 4545 CREEK RD |  | ETHICON INC. | Initial submission |  |  | N | Y | 45242 |  | US |  | Manufacturer report |  | 20130718 |  |  |  | CINCINNATI | N | 3170804 |  |  | 20130809 | 1 |  |  |  | 1, Required Intervention | 3299 | 151 |  |  | 513 | Y | Y |  | US | KATHY |  | 1 | 337 |  |  |  |  | N | Unknown | 1 | 3274209 |  |  |  | P.O. BOX 151, ROUTE 22 WEST | ETHICON SARL - NEUCHATEL | 2803 | RICE | B |  | 151 |  | N | HEALTH PROFESSIONAL | GYNECARE TVT RETROPUBIC SYSTEM | UNK |  | NJ | MESH, SURGICAL, SYNTHETIC, UROGYNECOLOGIC | ETHICON INC. | No | SOMERVILLE | 1 |  | UNK |  | 20130809 |  | P.O. BOX 151, ROUTE 22 WEST |  |  | UNK | 8876 | FTL |  | US | 3303434 | R
[22 empty] | IT WAS REPORTED THAT THE PATIENT UNDERWENT A GYNECOLOGICAL PROCEDURE IN 1999 AND A MESH, AMS SPARC, BARD ALIGN, BOSTON SCIENTIFIC OBTRYX AND COLOPLAST ARIS WERE IMPLANTED. IT WAS REPORTED THAT SHE EXPERIENCED PAIN, EROSION OF HER INTERNAL BODILY TISSUE AND OTHER INJURIES FOLLOWING THE PROCEDURE. IT WAS REPORTED THAT THE PATIENT HAS UNDERGONE MULTIPLE SURGERIES AND REVISIONARY PROCEDURES. NO ADDITIONAL INFORMATION WAS PROVIDED. | 20023738 | Description of Event or Problem | 1
3007566237-2013-03351 | MN | 2013-03-04 | Injury |  | Foreign | Company representation | 7000 CENTRAL AVENUE NE RCW215 |  | N | I | 1131 | TOLOCHENAZ | OTHER |  |  | MN |  |  | US |  | MINNEAPOLIS | IT WAS REPORTED THAT THERE WAS NOT GOOD TELEMETRY. IT WAS STATED THAT IT WAS PROBABLY DUE TO AN IMPLANTABLE NEUROSTIMULATOR (INS) DEPTH ISSUE. X-RAYS WERE TAKEN AND DID NOT SUGGEST THAT THE INS WAS FLIPPED. HOWEVER, IT WAS STATED THAT THE INS APPEARED TO BE TILTED. A REVISION WAS SCHEDULED FOR (B)(6)-2013. | 21037705 | Description of Event or Problem | 1 |  | 20131014 | 55421 |  | 20130923 | 7000 CENTRAL AVENUE NE RCW215 |  | IPG MFG SWITZERLAND |  | Initial submission | Followup | N | Y | 55432 |  | US |  | Manufacturer report |  | 20131115 |  |  |  | MINNEAPOLIS | Y | 3304308 | 55432 | MN | 20131112 | 1 |  |  |  | 1, Required Intervention | 9619 | 1200 |  |  | 763 | Y | Y |  | CH | CHAD |  | 1 | 514 |  |  |  |  | N | Unknown | 1 | 3407713 |  |  |  | ROUTE DU MOLLIAU 31 | MEDTRONIC NEUROMODULATION |  | HEDLUND | V |  | 1200 |  | N | HEALTH PROFESSIONAL | INTERSTIM |  | 20140828 | MN | STIMULATOR, ELECTRICAL, IMPLANTABLE, FOR INCONTINENCE | IPG MFG SWITZERLAND | Device was returned to manufacturer | TOLOCHENAZ | 1 | DA | 3023 |  | 20131014 |  | ROUTE DU MOLLIAU 31 | 1131 |  | 3023 | 55421 | EZW | 20131029 | CH | 3432718 | Y
[22 empty] | (B)(4). | 21037706 | Additional Manufacturer Narrative | 1
[22 empty] | ANALYSIS OF THE DEVICE, MODEL # 3023, SN (B)(4), FOUND NO SIGNIFICANT ANOMALIES. THE CAN WAS SEVERELY DENTED, BUT IT DID NOT AFFECT THE DEVICE FUNCTION. | 23461179 | Additional Manufacturer Narrative | 1
[22 empty] | ADDITIONAL REVIEW SHOWED THE CORRECT MANUFACTURING SITE WAS SITE # 9614453, AS THE DEVICE SERIAL NUMBER IS NOW KNOWN. (B)(4). ANALYSIS RESULTS WERE NOT AVAILABLE AS OF THE DATE OF THIS REPORT. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN ANALYSIS IS COMPLETE. | 31809724 | Additional Manufacturer Narrative | 1
[22 empty] | ADDITIONAL INFORMATION RECEIVED REPORTED THAT PRIOR TO THE REVISION; THE IMPLANTABLE NEUROSTIMULATOR (INS) WAS NOT PALPABLE UNDER OR IN DIRECT AREA OF THE INCISION, THE INS WAS PALPABLE LATERAL FROM THE INCISION, AND NO COMMUNICATION WAS POSSIBLE WITH THE PATIENT OR CLINICIAN PROGRAMMERS. DURING THE PROCEDURE, UNIPOLAR ELECTROCAUTERY AND A SURGICAL CLAMP WERE USED FOR REMOVAL. THE INS WAS REPLACED AND INTERROGATION WITH THE CLINICIAN PROGRAMMER WAS SUCCESSFUL AFTER INCISION CLOSURE. IT WAS STATED THAT THE CLAMP CAUSED DAMAGE TO THE OLD INS DURING REMOVAL. ONCE REMOVED, THE OLD INS WAS INTERROGATED AND TURNED OFF. THE PATIENT WAS DOING FINE WITH NEW SYSTEM AND THERE WAS NOW A DIRECT RESPONSE WITH THE INS. | 24039230 | Description of Event or Problem | 1
2411512-2014-00013 | NJ |  |  |  | Company representation | Other | 532 BROADHOLLOW ROAD |  | N | * |  | MELVILLE | PHARMACIST |  |  | NY |  | 20141014 | US | NA | MELVILLE | CASE RECEIVED FROM A PHARMACIST, RADIOLOGIST AND A RADIOLOGY TECHNICIAN VIA (B)(6) ON (B)(6) 2014 AND FORWARDED TO (B)(4) ON THE SAME DAY. ON 10/21/2014, LIVE FOLLOW-UP WAS RECEIVED FROM THE SAME SOURCES AND INCLUDED IN THE INITIAL REPORT. HEALTH PROFESSIONALS REPORT: A (B)(6) MALE PATIENT WITH MEDICAL HISTORY OF STROKE, APPENDECTOMY AT THE AGE OF (B)(6) AND ABLATION OF ONE MICRO POLYP DURING COLONOSCOPY IN (B)(6) 2014, BUT WITHOUT HISTORY OF GASTROINTESTINAL DISEASE, UNDERWENT A CT VIRTUAL COLONOSCOPY WITH PROTOCO21 FOR COLON INSUFFLATION SET FOLEY CATHETER (REF. 6450, LOT NUMBER UNK) ON (B)(6) 2014 DUE TO INCOMPLETE VISUALIZATION OF THE RIGHT ANGLE OF COLON DURING THE COLONOSCOPY IN (B)(6) 2014. THE BALLOON WAS INFLATED WITH 50 ML OF AIR TWICE (35 ML AND 15 ML). BEFORE INTRODUCTION INTO THE RECTUM THE BALLOON INFLATION HAD BEEN TESTED WITH THE SAME AMOUNT OF AIR. DURING THE PROCEDURE THE BALLOON EXPLODED WHILE THE PATIENT CHANGED FROM DORSAL TO LATERAL POSITION. THE PATIENT HAD NOTICED A CHANGE AND THE IMAGING TEAM REALIZED THAT THE BALLOON HAD EXPLODED. THE PROCEDURE WAS STOPPED, BUT COULD BE RESTARTED AND FINALIZED WITHOUT PROBLEM ON THE SAME DAY WITH ANOTHER BALLOON ADMINISTRATION SET FOLEY CATHETER. NO MEDICAL CONSEQUENCES AT ALL OCCURRED FOR THE PATIENT. OUTCOME: UNK. THIS CASE IS MEDICALLY CLOSED. NOVEMBER 21, 2014: FOLLOW-UP INFORMATION WAS RECEIVED FROM THE (B)(4)AUTHORITY (B)(6) ON (B)(6) 2014 AND FORWARDED TO (B)(4) ON THE SAME DAY. FURTHERMORE, A RECEIPT CONFIRMATION OF THE LAST SUBMISSION WITH THE ANSM NUMBER WAS RECEIVED VIA LANDMAIL DIRECTLY AT (B)(4). THE AUTHORITY CONFIRMED THE FINAL ASSESSMENT OF CASE NUMBER (B)(4), REGARDING AN INCIDENT ON ADMINISTRATION SET FOLEY CATHETER. NO FURTHER INVESTIGATION WAS PLANNED. OUTCOME: UNK. THIS CASE IS MEDICALLY CLOSED. (B)(4). | 29322080 | Description of Event or Problem | 1 |  | 20141211 | 11747 | BLDG H | 20141014 | 259 PROSPECT PLAINS RD. |  | BRACCO DIAGNOSTICS, INC. | Initial submission |  |  |  | Y | 8831 |  | US |  | Manufacturer report |  | 20141121 |  |  |  | MONROW TWP | N | 4256948 | 11747 | NY | 20141211 | 1 | KARDEGIC (ACETYLSALICYLATE LYSINE); |  |  |  | 5181 |  |  |  | 800 | N |  |  | US |  |  | 1 | 257 |  |  |  |  | Y | Unknown | 1 | 4360353 |  |  |  | 532 BROADHOLLOW RD. | EZEM |  |  | B | NA |  |  | N | 0 | ADMINISTRATION SET FOLEY CATHETER | NI |  | NY | PROTOCO2L COLON INSUFFLATOR ADMINISTRATION SET FOLEY CATHETER | BRACCO DIAGNOSTICS, INC. |  | MELVILLE | 1 |  | NI |  | 20141211 |  | 532 BROADHOLLOW RD. |  |  | 6470 | 11747 | FCX |  | US | 4411209 | *
[22 empty] | A (B)(6) MALE PATIENT UNDERWENT A CT COLONOSCOPY WITH PROTOCO2L SET FOR COLON INSUFFLATION, AS DURING A PREVIOUS COLONOSCOPY THE RIGHT COLON ANGLE COULD NOT BE VISUALIZED. THE PATIENT HAD NO HISTORY OF GASTROINTESTINAL DISEASE. HOWEVER, DURING THE COLONOSCOPY PERFORMED 3 MONTHS AGO, ONE MICRO POLYP WAS REMOVED. OTHER MEDICAL HISTORY INCLUDED APPENDECTOMY AT THE AGE OF (B)(6) AND STROKE. DURING THE PROCEDURE, WHILE CHANGING POSITION, THE PATIENT NOTICED A CHANGE, AND SUBSEQUENTLY THE IMAGING TEAM NOTED THAT THE BALLOON HAD EXPLODED. THERE WAS NO CLINICAL CONSEQUENCE FOR THE PATIENT DUE TO THE EVENT AND THE PROCEDURE COULD BE COMPLETED WITH ANOTHER BALLOON ON THE SAME DAY. TO PERFORM CT COLONOSCOPY THIN CATHETER WITH SMALL BALLOON IS GENERALLY USED. EXPLOSION OF THE BALLOON MIGHT OCCUR AND THIS COULD BE THE RESULT OF A BALLOON OVER-DISTENTION OR OF A VARIATION OF THE RESULTANT FORCE BETWEEN THE EXTERNAL AND INTERNAL PRESSURE OF THE BALLOON. IN THIS CASE, THE BALLOON INFLATION HAD BEEN TESTED BEFORE ITS INSERTION INTO THE RECTUM. THE TOTAL AMOUNT OF AIR USED TO INFLATE THE BALLOON WAS 50 ML (35 ML PLUS 15 ML). THE INCIDENT HAS BEEN CONSIDERED POSSIBLY RELATED TO THE PROCEDURE SINCE THE BALLOON EXPLOSION OCCURRED WHEN THE PATIENT CHANGE POSITION (E.G. FROM DORSAL DECUBITUS TO LATERAL DECUBITUS). IN FACT, DURING THE CT COLONOSCOPY, WHEN THE PATIENT CHANGE THE POSITION FROM SUPINE TO PRONE A HIGHER PRESSURE UP TO 40-50 MMHG COULD BE GENERATED INTO THE COLON INSTEAD OF THE OPTIMAL ONE OF 25 MMHG. IN THIS CIRCUMSTANCE, A VARIATION OF THE RESULTANT FORCE BETWEEN THE EXTERNAL AND INTERNAL PRESSURE OF THE BALLOON COULD OCCUR AND DETERMINE AND EXPLOSION/DEFLATION OF THE BALLOON. BASED ON THE ABOVE, EVEN THOUGH THE EVENT IS CONSIDERED NOT MANUFACTURING RELATED BUT PROCEDURAL RELATED, THE MANUFACTURER HAS DECIDED TO HAVE A CONSERVATIVE APPROACH AND TO REPORT THIS CASE. | 29322062 | Additional Manufacturer Narrative | 1
2032227-2014-49432 | CA | 2013-01-01 | Malfunction | Consumer |  |  | 18000 DEVONSHIRE STREET |  | N | U |  | NORTHRIDGE | OTHER |  |  | CA |  | 20141010 | US |  | NORTHRIDGE | CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT. THE DEVICE HAS BEEN RETURNED, BUT NOT YET EVALUATED. FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. | 27883847 | Additional Manufacturer Narrative | 1 |  | 20141107 | 91325 |  | 20141010 | 18000 DEVONSHIRE STREET |  | MEDTRONIC MINIMED |  | Initial submission | Followup | N | Y | 91325 | 1219 | US |  | Manufacturer report |  | 20141020 |  |  |  | NORTHRIDGE | Y | 4135069 | 91325 | CA | 20141107 | 1 |  |  |  |  | 4313 |  |  |  | 818 | N | N |  | US | CHRISTOPHER |  | 1 | 576 |  | MR. |  |  | N | Unknown | 1 | 4238497 |  |  |  | 18000 DEVONSHIRE STREET | MEDTRONIC MINIMED | 1219 | TUPPER | B |  |  |  | N | LAY USER/PATIENT | PARADIGM REAL-TIME REVEL INSULIN INFUSION PUMP |  |  | CA | INSULIN INFUSION PUMP / SENSOR AUGMENTED | MEDTRONIC MINIMED | Device was returned to manufacturer | NORTHRIDGE | 1 | DA | MMT-723LNAL |  | 20141107 |  | 18000 DEVONSHIRE STREET |  |  | MMT-723LNAL | 91325 | OYC | 20141020 | US | 4268770 | Y
[22 empty] | CUSTOMER REPORTED THAT THE INSULIN PUMP ALARMED BUTTON ERROR. BLOOD GLUCOSE VALUE WAS 61 MG/DL; TREATED BY DRINKING JUICE. NOTHING FURTHER REPORTED. | 27883845 | Description of Event or Problem | 1
[22 empty] | NO UNEXPECTED BUTTON ERROR ALARM OR AUDIO ANOMALY WAS NOTED. ALL OF THE BUTTONS FUNCTIONED PROPERLY. HOWEVER, CORRODED KEYPAD TRACES WERE NOTED DURING THE VISUAL INSPECTION. THE INSULIN PUMP HAD A CRACKED DISPLAY WINDOW, A CRACKED CASE AT THE DISPLAY WINDOW CORNERS, CRACKED RESERVOIR TUBE LIP, SCRATCHED RESERVOIR TUBE WINDOW, CRACKED BATTERY TUBE THREADS AND A MISSING END CAP STICKER. | 29745444 | Additional Manufacturer Narrative | 1
1061932-2014-01833 | CA | 2011-04-01 | Malfunction |  | Health Professional | User facility | 11800 SW 147TH AVENUE |  | N | U |  | BREA | OTHER HEALTH CARE PROFESSIONAL |  |  | CA |  | 20140714 | US | NA | MIAMI | THE CUSTOMER REPORTED THE UNICEL DXH 800 COULTER CELLULAR ANALYSIS SYSTEM GENERATED INCOMPLETE RESULTS FOR THE DIFFERENTIAL (DIFF) AND NUCLEATED RED BLOOD CELL (NRBC). THE CUSTOMER ALSO REPORTED THE INSTRUMENT GENERATED 'LOW EVENT RATE' ERRORS. THERE WAS NO REPORTED EXPOSURE TO MUCOUS MEMBRANES OR OPEN WOUNDS. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF GLOVES AND A LAB COAT AT THE TIME OF THE INCIDENT. THERE WERE NO ERRONEOUS RESULTS REPORTED OUT OF THE LABORATORY. THERE WAS NO DEATH, INJURY, OR TO PATIENT TREATMENT ATTRIBUTED TO THIS EVENT. | 25978435 | Description of Event or Problem | 1 |  | 20140806 | 92821 |  | 20140714 | 250 S. KRAEMER BLVD. |  | BECKMAN COULTER | Initial submission |  |  | N | Y | 92821 |  | US |  | Manufacturer report |  | 20140714 |  |  |  | BREA | Y | 3868720 | 33196 | FL | 20140806 | 1 |  |  |  |  | 4941 |  |  |  | 714 | N | N |  | US | DUNG |  | 1 | 961 |  | MS. |  |  | Y | Unknown | 1 | 3982401 |  |  |  | 250 S. KRAEMER BLVD. | BECKMAN COULTER |  | NGUYEN | B | SW VERSION: 2.0 |  |  | N | HEALTH PROFESSIONAL | UNICEL DXH 800 COULTER CELLULAR ANALYSIS SYSTEM | NA |  | CA | COUNTER, DIFFERENTIAL CELL | BECKMAN COULTER | Yes | BREA | 1 |  | 629029 |  | 20140806 |  | 250 S. KRAEMER BLVD. |  |  | NA | 92821 | GKZ |  | US | 4008574 | Y
[22 empty] | A BECKMAN COULTER (BEC) FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED TO ASSESS THE INSTRUMENT. THE FSE CONFIRMED THE SAMPLE LINE (TUBING) HAD BECOME DISCONNECTED FROM THE BOTTOM OF THE FLOWCELL CAUSING A SMALL CONTAINED LEAK. THE VOLUME OF THE LEAK WAS UNKNOWN. THE FSE REPLACED THE SAMPLE LINE (TUBING), RESOLVING THE LEAK, INCOMPLETE RESULTS, AND ERRORS. THE FSE ALSO BLEACHED THE FLOWCELL WASTE PATHWAY AS PART OF TROUBLESHOOTING. WHILE ONSITE, THE FSE OBSERVED THE LEAK HAD DAMAGED THE MEDIAN ANGLE LIGHT SCATTER (MALS) SENSOR CAUSING ERRATIC OFFSETS. THE FSE REPLACED THE MALS SENSOR, RESOLVING THE ISSUE. (B)(4). | 25978476 | Additional Manufacturer Narrative | 1
2939301-2014-20356 |  |  | Malfunction |  | Foreign | Consumer | GUBELSTRASSE 34 |  | N | I | 6300 | ZUG | PATIENT |  | 20140804 |  |  |  | SZ |  | ZUG | ON (B)(6) 2014, THE REPORTER CONTACTED LIFESCAN (B)(4), ALLEGING DATA PORT ISSUE. THIS COMPLAINT IS BEING REPORTED BECAUSE THE REPORTED ISSUE WAS NOT RESOLVED WITH TROUBLESHOOTING. THERE WAS NO INDICATION THAT THE PRODUCT CAUSED OR CONTRIBUTED TO AN ADVERSE EVENT. | 25949073 | Description of Event or Problem | 1 |  | 20140811 |  |  | 20140804 | GUBELSTRASSE 34 |  | LIFESCAN EUROPE, A DIVISION OF CILAG GMBH INTL | Initial submission |  |  | N | Y |  |  | SZ |  | Manufacturer report |  | 20140804 |  |  | 6300 | ZUG | Y | 3894312 |  |  | 20140811 | 1 | [7 empty] | 6300 |  | N | N |  | SZ | M. CHIUSANO / |  | 1 | [5 empty] | N | Unknown | 1 | 3997697 | Y |  |  | GUBELSTRASSE 34 | LIFESCAN EUROPE, A DIVISION OF CILAG GMBH INTL |  | Y. WANDOLSKI | B |  |  |  | N | LAY USER/PATIENT | OT VERIO METER | 3519180 |  |  | GLUCOSE MONITORING SYS/KIT | LIFESCAN EUROPE, A DIVISION OF CILAG GMBH INTL | No | ZUG | 1 |  |  |  | 20140811 |  | GUBELSTRASSE 34 | 6300 |  |  |  | NBW |  | SZ | 4024649 | R
2531779-2014-22068 | PA | 2012-10-18 | Malfunction | Consumer |  |  | 200 LAWRENCE DR |  | N | U |  | WEST CHESTER | PATIENT |  |  | PA |  |  | US |  | WEST CHESTER | ON (B)(6) 2014, THE REPORTER CONTACTED ANIMAS, ALLEGING A CASING/CONDITION (MOISTURE INGRESS) ISSUE. THIS COMPLAINT IS BEING REPORTED BECAUSE THE REPORTED ISSUE WAS NOT RESOLVED WITH TROUBLESHOOTING. THERE WAS NO INDICATION THAT THE PRODUCT CAUSED OR CONTRIBUTED TO AN ADVERSE EVENT. | 25838183 | Description of Event or Problem | 1 |  | 20140801 | 19380 |  | 20140723 | 200 LAWRENCE DR |  | ANIMAS CORPORATION |  | Followup | Initial submission | N | Y | 19380 | 3428 | US |  | Manufacturer report |  | 20140723 |  |  |  | WEST CHESTER | Y | 3869558 | 19380 | PA | 20140919 | 1 |  |  |  |  | 1808 | 3428 |  |  | 484 | N | N |  | US | KARIN |  | 1 | 356 |  |  |  |  | N | Unknown | 1 | 3972961 | Y |  |  | 200 LAWRENCE DR | ANIMAS CORPORATION | 3428 | SARGRAD | B | 1-MCBC-64 | 3428 |  | N | LAY USER/PATIENT | ONETOUCHPING GLUCOSEMGMTSYSTEM |  |  | PA | INSULIN INFUSION PUMP | ANIMAS CORPORATION | Device was returned to manufacturer | WEST CHESTER | 1 | 21 MO |  |  | 20140801 |  | 200 LAWRENCE DR |  |  |  | 19380 | LZG | 20140813 | US | 3999559 | Y
[22 empty] | THE PUMP HAS NOT BEEN RETURNED TO ANIMAS FOR EVALUATION. IF THE DEVICE IS RETURNED, AN EVALUATION SHALL BE COMPLETED AND A SUPPLEMENTAL REPORT WILL BE FILED. NO CONCLUSIONS CAN BE MADE AT THIS TIME. | 25838169 | Additional Manufacturer Narrative | 1
[22 empty] | FOLLOW-UP #1: DATE OF SUBMISSION 09/19/2014. DEVICE EVALUATION: THE DEVICE HAS BEEN RETURNED AND EVALUATED BY PRODUCT ANALYSIS ON (B)(4) 2014 WITH THE FOLLOWING FINDINGS: THE DISPLAY LENS LEAK WITH EVIDENCE OF MOISTURE CORROSION ON DISPLAY BOARD. THERE WAS MOISTURE CORROSION NEXT TO U 38. THE DISPLAY IS DIM AND PINK. | 28087364 | Additional Manufacturer Narrative | 1
1225058-2014-11326 | SC | 2002-03-01 | Malfunction | Company representation |  |  |  |  | N | * |  | MISQUAMICUT | OTHER HEALTH CARE PROFESSIONAL |  |  | RI |  | 20141113 |  |  |  | AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. WITHIN THIS OCCURRENCE ARE 825 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER/S (MRN) IDENTIFIED. | 30969817 | Description of Event or Problem | 1 |  | 20141212 | 2891 |  | 20141212 | 10 SOUTH ACADEMY ST. |  | AGFA HEALTHCARE CORPORATION | Initial submission |  |  | N | Y | 29601 |  | US |  | Manufacturer report |  | 20141113 | Other |  |  | GREENVILLE | Y | 4461290 |  |  | 20141212 | 1 |  |  |  |  | 1754 |  |  |  | 864 | N |  |  | US | DEBORAH |  | 1 | 421 |  | MRS. |  |  | Y | No | 1 | 4581985 |  |  |  | 1 CROSSWIND RD. |  |  | HUFF | B | NA |  |  | N | 0 | IMPAX CV OUTBOUND (RESULTS MANAGEMENT 2.08) | NA |  | RI | PICTURE ARCHIVING AND COMMUNICATION | AGFA HEALTHCARE CORPORATION | Yes | MISQUAMICUT | 1 |  | NA |  | 20141212 |  | 1 CROSSWIND RD. |  |  | IMPAX CV 2.08 | 2891 | LLZ |  | US | 4688933 | Y
[22 empty] | AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATIONS HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: ORIGINAL SENTENCE OF: "THERE IS MINIMAL AORTIC STENOSIS PRESENT" HAS BEEN CHANGED TO: "THERE IS AORTIC VALVE SCLEROSIS WITHOUT SIGNIFICANT AORTIC VALVE STENOSIS". IN THIS ISSUE, THE ORIGINAL SENTENCE INDICATES THE PRESENCE OF MINIMAL AORTIC STENOSIS. THE CHANGED SENTENCE CONTAINS CONTRADICTORY CLINICAL INFO AND STATES THAT THERE IS AV SCLEROSIS WITHOUT SIGNIFICANT AV STENOSIS. THIS CHANGE WOULD NOT REFLECT THE INTENT OF THE ORIGINAL READING CARDIOLOGIST AND WOULD BE INACCURATE. POTENTIAL IMPACT: IF THE CHANGE WAS MADE AFTER REPORTS HAD BEEN GENERATED WITH THE ORIGINAL SENTENCE SELECTED, THOSE REPORTS PRIOR TO THE CHANGE COULD DISPLAY INCORRECT CLINICAL DATA IF VIEWED IN RM OF IF USED AS A TEMPLATE FOR A REPEAT STUDY ON THAT PT. THERE HAS BEEN NO REPORTED PT HARM FOR THIS OCCURRENCE. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA VIA REFERENCE #Z-2112-10. AGFA WILL FOLLOW THE RMAT POST MARKET VERIFICATION WORK INSTRUCTIONS TO CORRECT THE SENTENCE FINDING. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(6). | 30969876 | Additional Manufacturer Narrative | 1
2183613-2013-00717 | MN | 1997-08-29 | Malfunction |  | Company representation | Health Professional | 8200 CORAL SEA ST NE |  | N | R |  | MILACA | HEALTH PROFESSIONAL |  |  | MN |  | 20130429 | US |  | MOUNDS VIEW | IT WAS REPORTED THAT THE VENTRICULAR OUTPUT ON THE EXTERNAL PULSE GENERATOR (EPG) WOULD NOT ADJUST FROM 10 MILLIAMPS (MA). THE ATRIAL CHANNEL WAS GOOD AT 5 MA, HOWEVER IT WAS ONLY DELIVERING AT 10 MA OUTPUT WHEN SET TO 20 MA. IT WAS ALSO REPORTED THE VENTRICULAR OUTPUT WILL ONLY FLUCTUATE FROM 10 +/- 1 MA WHEN DIAL IS TURNED EITHER DIRECTION. THE EPG WAS RETURNED FOR REPAIR. THERE WAS NO PATIENT INVOLVEMENT. | 21465004 | Description of Event or Problem | 1 |  | 20130608 | 56353 |  | 20130429 | 8200 CORAL SEA ST NE |  | MEDTRONIC MILACA, INC. |  | Followup | Initial submission | N | Y | 55112 |  | US |  | Manufacturer report |  | 20130618 |  |  |  | MOUNDS VIEW | Y | 3051512 | 55112 | MN | 20130810 | 1 |  |  |  |  | 583 | 3728 |  |  | 763 | N | N |  | US | NASHOANE |  | 1 | 526 |  |  |  |  | Y | Unknown | 1 | 3154977 |  |  |  | 900 SIXTH AVENUE NE | MEDTRONIC CARDIAC RHYTHM DISEASE MGMT |  | FULWOOD-KELLEY | B |  | 3728 |  | N | HEALTH PROFESSIONAL |  |  |  | MN | PULSE-GENERATOR, DUAL CHAMBER, IMPLANTABLE | MEDTRONIC MILACA, INC. | Device was returned to manufacturer | MILACA | 1 | DA | 5388 |  | 20130608 |  | 900 SIXTH AVENUE NE |  |  | 5388 | 56353 | LWP | 20130506 | US | 3184210 | Y
[22 empty] | PRODUCT EVENT SUMMARY: ANALYSIS CONFIRMED THE CUSTOMER COMMENT THAT THE VENTRICLE OUTPUT COULD NOT BE ADJUSTED AND IT WAS ATTRIBUTED TO THE ENCODER FLEX BEING DAMAGED. ANALYSIS ALSO FOUND THE UPPER AND LOWER CASES BROKEN, AND THE LOWER CASE, BATTERY RELEASE, RING COVER, SIDE BAIL COVERS, LEAD FLEX COVER, HEART BLOCK, HEART WIRE CONTACTS, BATTERY DRAWER AND HEART LEAD FLEX WERE CONTAMINATED, THE BATTERY CONTACTS WERE COMPRESSED AND CONTAMINATED, THE BATTERY DRAWER O-RING WAS MISSING, THE BATTERY FLEX WAS CONTAMINATED AND CORRODED, AND THAT A LIQUID CRYSTAL DISPLAY (LCD) SCREW WAS BEING USED AS A MISSING CASE SCREW. | 21465023 | Additional Manufacturer Narrative | 1
[22 empty] | IT WAS REPORTED THAT THE VENTRICULAR OUTPUT ON THE EXTERNAL PULSE GENERATOR (EPG) WOULD NOT ADJUST FROM 10 MILLIAMPS (MA). THE ATRIAL CHANNEL WAS GOOD AT 5 MA, HOWEVER IT WAS ONLY DELIVERING AT 10 MA OUTPUT WHEN SET TO 20 MA. IT WAS ALSO REPORTED THE VENTRICULAR OUTPUT WILL ONLY FLUCTUATE FROM 10 +/- 1 MA WHEN DIAL IS TURNED EITHER DIRECTION. THE EPG WAS RETURNED FOR REPAIR. THERE WAS NO PATIENT INVOLVEMENT. | 20325178 | Description of Event or Problem | 1
[22 empty] | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. (B)(4). | 20325180 | Additional Manufacturer Narrative | 1
9680959-2014-00320 | UT | 2011-02-01 | Malfunction |  | User facility | Company representation | 14 WILHELM-MAISEL-STR. |  | N | R |  | WENDELSTEIN | OTHER |  |  |  |  | 20140127 | GM |  | WENDELSTEIN | THE CUSTOMER REPORTED THE IMAGE DISAPPEARED SUDDENLY. IT IS LIKELY THIS SYSTEM INTERMITTENTLY FAILED TO DISPLAY A LIVE IMAGE. THERE ARE NO REPORTS OF PATIENT INJURY OR DEATH ASSOCIATED WITH THIS EVENT. | 24811841 | Description of Event or Problem | 1 |  | 20140227 |  |  | 20140307 | 384 WRIGHT BROTHERS DR |  | GE OEC MEDICAL SYSTEMS GMBH | Initial submission |  |  | N | Y | 84116 |  | US |  | Manufacturer report |  | 20140225 |  |  |  | SALT LAKE CITY | Y | 3725868 |  |  | 20140227 | 1 |  |  |  |  | 4952 |  |  | D-90530 | 801 | N | N |  | GM |  |  | 1 | 536 |  |  |  |  | Y | Unknown | 1 | 3829313 |  |  | 20140307 | 14 WILHELM-MAISEL-STR. | GE OEC MEDICAL SYSTEMS GMBH |  |  | B |  |  |  | N | HEALTH PROFESSIONAL | 7900 |  |  |  | FLUOROSCOPIC X-RAY | GE OEC MEDICAL SYSTEMS GMBH | Yes | WENDELSTEIN | 1 | DA |  |  | 20140227 |  | 14 WILHELM-MAISEL-STR. |  |  | 7900 |  | JAA |  | GM | 3889452 | Y
[22 empty] | A GE SERVICE REPRESENTATIVE PERFORMED AN ON SITE INVESTIGATION. THE REPORTED ISSUE COULD NOT BE DUPLICATED. THE BOARDS AND CONNECTORS WERE RESEATED DURING THE SERVICE CALL. THE SYSTEM WAS TESTED AND FOUND TO BE WORKING AS INTENDED AND PUT BACK INTO SERVICE. | 24811854 | Additional Manufacturer Narrative | 1
2028159-2013-00841 | TX | 2001-06-01 | Malfunction | Health Professional |  |  | 15800 ALTON PKWY. |  | N | R |  | IRVINE | OTHER |  |  | CA |  | 20130410 | US | NA | IRVINE | THE COMPANY REP EXAMINED THE SYSTEM AND COULD NOT DUPLICATE THE POWER DOWN ISSUE REPORTED. THE COMPANY REP REPLACED A HALOGEN LAMP THAT WAS OUT, WHICH IS NOT RELATED TO THE REPORTED EVENT. IT WAS NOTED THAT THE SYSTEM WAS PLUGGED INTO A POWER STRIP BEFORE THE SYSTEM SHUT DOWN. ACCORDING TO THE SYSTEM OPERATOR'S MANUAL, THE SYSTEM POWER CORD IS A MEDICAL GRADE POWER CORD WITH THE LEAST LEAKAGE CURRENT PER FOOT RATING AVAILABLE. EXTENSION OF THE POWER CORD BY HOSPITAL STAFF IS NOT RECOMMENDED. THE SYSTEM WAS THEN TESTED AND MET ALL PRODUCT SPECIFICATIONS. A REVIEW OF COMPLAINTS FOR THE LAST (B)(4) DID NOT INDICATE ANY ADD'L SIMILAR REPORTS FOR THIS SYSTEM. THE ROOT CAUSE CANNOT BE DETERMINED CONCLUSIVELY. (B)(4). | 18821414 | Additional Manufacturer Narrative | 1 |  | 20130503 | 92618 | TC-40 | 20130410 | 6201 SOUTH FREEWAY |  | ALCON MANUFACTURING, LTD. | Initial submission |  |  | N | Y | 76134 | 3818 | US |  | Manufacturer report |  | 20130410 |  |  |  | FORT WORTH | Y | 2998700 | 92618 | CA | 20130503 | 1 | UNK; |  |  |  | 2442 | 3818 |  |  | 817 | N | N |  | US | ARTHUR |  | 1 | 615 |  |  |  |  | Y | Unknown | 1 | 3102145 |  |  |  | 15800 ALTON PKWY. | ALCON MANUFACTURING, LTD. | 3818 | CIOCIOLA | B | 4.52 | 3818 |  | N | HEALTH PROFESSIONAL | ACCURUS 800CS | NA |  | CA | PHACOFRAGMENTATION SYSTEM | ALCON MANUFACTURING, LTD. | Yes | IRVINE | 1 | DA | 8065741008 |  | 20130503 |  | 15800 ALTON PKWY. |  |  | ACCURUS | 92618 | HQC |  | US | 3146599 | Y
[22 empty] | A CUSTOMER REPORTED THE SYSTEM'S FAN WAS MAKING A NOISE AND THE SYSTEM SHUT DOWN ON ITS OWN DURING A PROCEDURE. THE PROCEDURE WAS COMPLETED USING THE SAME SYSTEM AND ACCESSORIES. THERE WAS NO IMPACT TO THE PATIENT. | 18821410 | Description of Event or Problem | 1
3008973940-2013-00230 | MN | 2013-09-03 | Injury |  | Health Professional | Company representation | 8200 CORAL SEA ST NE |  | N | I |  | SINGAPORE | PHYSICIAN |  |  |  |  | 20131018 | US |  | MOUNDS VIEW | IT WAS REPORTED THAT ONE DAY POST-IMPLANT, THE LEAD WAS OBSERVED WITH RISEN THRESHOLD THAT MEASURED AT A HIGH VALUE, AND THE LEAD WAS UNABLE TO PACE EVEN AT HIGH THRESHOLD VALUES. THE LEAD WAS EXPLANTED AND WAS REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. | 21918850 | Description of Event or Problem | 1 |  | 20131210 |  |  | 20131024 | 8200 CORAL SEA ST NE |  | MEDTRONIC SINGAPORE OPERATIONS |  | Initial submission | Followup | N | Y | 55112 |  | US |  | Manufacturer report |  | 20140117 |  |  |  | MOUNDS VIEW | Y | 3405984 | 55112 | MN | 20131210 | 1 |  | 1, Hospitalization | 1, Hospitalization, 2, Required Intervention |  | 583 |  |  |  | 763 | Y | Y |  | SG | NASHOANE |  | 1 | 526 |  |  |  |  | Y | Unknown | 1 | 3509409 |  |  |  | NASACO TECH CENTRE, 49 CHANGI | MEDTRONIC CARDIAC RHYTHM DISEASE MGMT |  | FULWOOD-KELLEY | V |  |  |  | N | LAY USER/PATIENT | CAPSURE SENSE |  | 20150827 |  | ELECTRODE, PACEMAKER, PERMANENT | MEDTRONIC SINGAPORE OPERATIONS | Device was returned to manufacturer | SINGAPORE | 1 | DA | 407458 |  | 20131210 |  | NASACO TECH CENTRE, 49 CHANGI |  |  | 407458 |  | DTB | 20131031 | SG | 3533779 | N
[22 empty] | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. (B)(4). | 21918855 | Additional Manufacturer Narrative | 1
[22 empty] | PRODUCT EVENT SUMMARY: THE FULL LEAD WAS RETURNED AND ANALYZED. ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. | 23991184 | Additional Manufacturer Narrative | 1
3004753838-2014-34130 | CA | 2014-08-19 | Malfunction | Consumer |  |  | 6340 SEQUENCE DRIVE |  | N | I |  | SAN DIEGO | OTHER |  |  | CA |  | 20141101 | US |  | SAN DIEGO | (B)(4). | 28024035 | Additional Manufacturer Narrative | 1 |  | 20141126 | 92121 |  | 20141101 | 6340 SEQUENCE DRIVE |  | DEXCOM, INC. | Initial submission |  |  | N | Y | 92121 |  | US |  | Manufacturer report |  | 20141101 |  |  |  | SAN DIEGO | Y | 4178464 | 92121 | CA | 20141126 | 1 |  |  |  |  | 200 |  |  |  | 800 | N | Y |  | US |  |  | 1 | 200 |  |  |  |  | N | No answer provided | 1 | 4281889 |  |  |  | 6340 SEQUENCE DRIVE | DEXCOM, INC |  |  | B | DI# 00386270000019 |  |  | N | LAY USER/PATIENT | G4 PLATINUM CONTINUOUS GLUCOSE MONITORING SYSTEM | 5190598 | 20150819 | CA | MDS | DEXCOM, INC. | No | SAN DIEGO | 1 |  |  |  | 20141126 |  | 6340 SEQUENCE DRIVE |  |  | 9500-27 | 92121 | MDS |  | US | 4311696 | R
[22 empty] | PATIENT CONTACTED DEXCOM TECHNICAL SUPPORT ON (B)(6) 2014, TO REPORT AN INACCURACY BETWEEN THE CONTINUOUS GLUCOSE MONITORING (CGM) SYSTEM AND THE BLOOD GLUCOSE (BG) METER ON (B)(6) 2014. PATIENT ENTERED THE BG METER VALUE TO CALIBRATE THE CGM DURING RAPID RATE OF CHANGE AGAINST THE USER'S GUIDE. THE PATIENT DID NOT REPORT INJURY OR MEDICAL INTERVENTION. | 28024034 | Description of Event or Problem | 1
2649622-2013-08749 | MN | 2012-10-23 | Injury |  | Company representation | Study | 8200 CORAL SEA ST NE |  | N | I |  | VILLALBA | PHYSICIAN |  |  | PR |  | 20130121 | US |  | MOUNDS VIEW | IT WAS REPORTED THAT THE DEVICE AND BOTH THE ATRIAL LEAD AND THE RIGHT VENTRICULAR (RV) LEAD WERE TWISTED AND ROTATED. THE RV LEAD DISLODGED AND WAS REPOSITIONED. THE ATRIAL LEAD AND DEVICE WERE REPOSITIONED. BOTH LEADS AND THE DEVICE ARE STILL IN USE. THE PATIENT IS ENROLLED IN THE SYSTEMS LONGEVITY STUDY. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. | 20066056 | Description of Event or Problem | 1 |  | 20130809 | 766 |  | 20130507 | 8200 CORAL SEA ST NE |  | MPRI | Initial submission |  |  | N | Y | 55112 |  | US |  | Manufacturer report |  | 20130507 |  |  |  | MOUNDS VIEW | Y | 3173748 | 55112 | MN | 20130809 | 1 | (B)(4) IMPLANTABLE PACING LEAD; | 1, Hospitalization | 1, Hospitalization, 2, Required Intervention |  | 583 |  |  |  | 763 | Y | Y |  | US | NASHOANE |  | 1 | 526 |  |  |  |  | Y | Unknown | 1 | 3277153 |  |  |  | ROAD 149 KM 56.3 | MEDTRONIC CARDIAC RHYTHM DISEASE MGMT |  | FULWOOD-KELLEY | B |  |  |  | Y | LAY USER/PATIENT | SPRINT QUATTRO SECURE |  | 20140710 | PR | DRUG-ELUTING PERMANENT DEFIBRILLATOR ELECTRODES | MPRI | No | VILLALBA | 1 |  | 6947M55 |  | 20130809 |  | ROAD 149 KM 56.3 |  |  | 6947M55 | 766 | NVY |  | US | 3306690 | *
[22 empty] | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. CONCOMITANT PRODUCTS: (B)(4) IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), 2013 (B)(6). (B)(4). | 20066060 | Additional Manufacturer Narrative | 1
2955842-2014-01734 | CA | 2013-11-01 | Malfunction | Other |  |  | 950 KIFER RD |  | N | U |  | SUNNYVALE | OTHER |  |  | CA | HOSPITAL | 20140224 |  |  | SUNNYVALE | THE INSTRUMENT HAS NOT BEEN RETURNED FOR EVALUATION; THEREFORE, THE ROOT CAUSE OF THE CUSTOMER REPORTED FAILURE MODE CANNOT BE DETERMINED. A FOLLOW-UP MDR WILL BE SUBMITTED IF THE INSTRUMENT IS RETURNED (POST FAILURE ANALYSIS EVALUATION) OR IF ADDITIONAL INFORMATION IS RECEIVED. THE CUSTOMER REPORTED COMPLAINT DOES NOT ITSELF CONSTITUTE A MDR REPORTABLE EVENT; HOWEVER; THE REPORTED MALFUNCTION IF TO RECUR COULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT. | 24433614 | Additional Manufacturer Narrative | 1 |  | 20140324 |  |  | 20140224 | 950 KIFER RD |  | INTUITIVE SURGICAL,INC. | Initial submission |  |  | N | Y |  |  |  |  | Manufacturer report |  | 20140224 |  |  | 94086 | SUNNYVALE | Y | 3593176 |  | CA | 20140324 | 1 | NA; | [6 empty] | 94086 |  | N | N |  |  | TABITHA |  | 1 | [5 empty] | N | No answer provided | 1 | 3696641 |  |  |  |  | INTUITIVE SURGICAL, INC., |  | REED | B |  |  |  | N | HEALTH PROFESSIONAL | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | M10131121 778 |  | CA | ENDOSCOPIC ELECTROSURGICAL INSTRUMENT | INTUITIVE SURGICAL,INC. | No | SUNNYVALE | 1 |  |  |  | 20140324 |  |  |  |  | 420205-05 |  | NAY |  |  | 3719306 | R
[22 empty] | IT WAS REPORTED THAT DURING THE REPROCESSING OF A FENESTRATED BIPOLAR FORCEPS INSTRUMENT, A BROKEN CABLE WAS IDENTIFIED. NOTHING REPORTEDLY FELL INTO A PATIENT AND NO PATIENT HARM, ADVERSE OUTCOME OR INJURY WAS REPORTED. | 24433635 | Description of Event or Problem | 1
2032227-2014-64015 | CA |  | Malfunction |  | Foreign | Company representation | 18000 DEVONSHIRE STREET |  | N | * |  | JUNCOS | OTHER |  |  | PR |  | 20141113 | US |  | NORTHRIDGE | IT WAS REPORTED THAT THE INSULIN PUMP WAS FILLING THE TUBING BY ITSELF WITH NO BUTTON HAVING BEEN PUSHED. NO EVENTS LED UP TO THIS. THE CUSTOMER'S BLOOD GLUCOSE WAS 7.8 MMOL/L. IT WAS ADVISED TO DISCONTINUE USE AND REVERT TO A BACK-UP PLAN. NOTHING FURTHER WAS REPORTED. | 28559948 | Description of Event or Problem | 1 |  | 20141211 | 777 |  | 20141113 | 18000 DEVONSHIRE STREET |  | MEDTRONIC PUERTO RICO OPERATIONS CO. |  | Initial submission | Followup | N | Y | 91325 | 1219 | US |  | Manufacturer report |  | 20141113 |  |  |  | NORTHRIDGE | Y | 4218588 | 91325 | CA | 20141211 | 1 |  |  |  |  | 4313 | 3869 |  |  | 818 | N | N |  | US | CHRISTOPHER |  | 1 | 576 |  | MR. |  |  | N | Unknown | 1 | 4322017 |  | ROAD 31 KM 24.4 |  | CEIBA NORTE IND. PARK #50 | MEDTRONIC MINIMED | 1219 | TUPPER | B |  | 3869 |  | N | 0 | PARADIGM INSULIN INFUSION PUMP |  |  | PR | INSULIN INFUSION PUMP / SENSOR AUGMENTED | MEDTRONIC PUERTO RICO OPERATIONS CO. | Device was returned to manufacturer | JUNCOS | 1 | DA | MMT-715WWL |  | 20141211 | ROAD 31 KM 24.4 | CEIBA NORTE IND. PARK #50 |  |  | MMT-715WWL | 777 | OYC | 20141222 | US | 4350440 | Y
[22 empty] | CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. | 28559926 | Additional Manufacturer Narrative | 1
[22 empty] | A COMPLETE ANALYSIS AND TESTING OF THE INSULIN PUMP SHOWED THAT IT WAS FUNCTIONING PROPERLY AND PASSED ALL FUNCTIONAL TESTING. AFTER TESTING, IT WAS CONCLUDED THAT THE DEVICE OPERATED WITHIN SPECIFICATIONS. | 30627178 | Additional Manufacturer Narrative | 1