## device_enforcement_limit25
recall_number | reason_for_recall | status | distribution_pattern | product_quantity | recall_initiation_date | state | event_id | product_type | product_description | country | city | recalling_firm | report_date | @epoch | voluntary_mandated | classification | code_info | @id | initial_firm_notification
Z-2338-2012 | Advanced Sterilization Products (ASP) is recalling (removing) the Sterrad 100S and Sterrad NX System cassettes because the camera used to verify printing on the Sterrad System cassettes was inadvertently disabled. The STERRAD 100S and STERRAD NX System cases were shipped without verification of proper printing. | Ongoing | Nationwide Distribution. | 7,550 units | 20120619 | CA | 62925 | Devices | STERRAD NX Cassette, Batch # 12A078. The STERRAD 100S and STERRAD NX sterilizations systems are low-temperature general purpose sterilizers used to sterilize heat and moisture sensitive reusable medical devices. The sterilant used is hydrogen peroxide supplied in cassette form as a separate accessory. The STERRAD Cassettes are used in conjunction with the STERRAD Sterilizers, and each cassette model is for use only in the matching sterilizer model. | US | Irvine | Advanced Sterilization Products | 20120919 | 1424553174.8365 | Voluntary: Firm Initiated | Class III | Product Code: 10133 | 00261c9b61c8635a201557251ea5d95a9342cc322b786859085100f9b4c901db | Letter
Z-2395-2012 | Embedded iron oxide glass defect which may have the potential to break off into the solution. | Ongoing | Worldwide Distribution-USA (nationwide) and the countries of Saudi Arabia, Canada and Barbados. | 191,094 units | 20120828 | IL | 63078 | Devices | Empty Evacuated Container 250 mL and 1000 mL. Empty Evacuated containers are used for mixing compatible medications or blood collections. | US | Lake Forest | Hospira Inc. | 20120926 | 1424553174.8365 | Voluntary: Firm Initiated | Class II | 250 mL size - Lot 11179DW, Exp, Date 1NOV2013, 11195DW, Exp. Date 1NOV2013 and 93854DW, Exp. Date 1SEP2012 - 1000 mL size - Lot 92835DW, Exp. Date 1AUG2012, Lot 96090DW, Exp. Date 1DEC2012, and Lot 96098DW, Exp. Date 1DEC2012 | 00325a3d5210ff60202b2f881fda1ae24d5c09daba758af3db7b95c26fe71a1b | Letter
Z-0347-2014 | Riverpoint Medical is recalling PGLA Surgical Suture (Vilet) because the label indicates the suture is undyed and it is actually a violet colored suture. | Terminated | US Distribution in the states of: AZ, FL, and TX. | 1932 pieces | 20121213 | OR | 65584 | Devices | Polyglycolic-Lactic Acid Absorbable Surgical Suture Label reads in part "Coated VILET *** (Polyglacin 910) Suture UNDYED BRANDED V423 *** Riverpoint ***" PGLA surgical suture is indicated for use in general soft tissue approximation and/or ligation, including use in opthalmic procedures, but not for use in cardiovascular and neurological tissues | US | Portland | Riverpoint Medical, LLC | 20131127 | 1424553174.8365 | Voluntary: Firm Initiated | Class II | Lot 120716-01 | 00374ab2fef30f1b646b062e520bf13dde94925d359cf93102bf521c890f6f71 | Letter
Z-1854-2012 | The recall was initiated by Accumetrics because of a field correction to the VerifyNow System; specifically the results reported when running the VerifyNow P2Y12 Test. Currently the VerifyNow P2Y12 Test reports three values: PRU, BASE and % Inhibition. The change implemented is the elimination of the BASE and % Inhibition results. There will be no change to the reporting of the PRU result which will be the only value reported. | Ongoing | US Nationwide Distribution | 1,167 units total | 20120206 | CA | 61761 | Devices | VerifyNow System, part #85005-6H Product Usage: The VerifyNow P2Y12 assay device is a whole blood assay used in the laboratory or point of care setting to measure the level of platelet P2Y12 receptor blockade. The VerifyNow System s intended for use with human whole blood and verify now test devices. The system should be operated by healthcare professionals trained on use of the system and in accordance with institution and policies and procedures. Accumetrics respresentatives will assist your institution in the installation of the system and operations of the operator. | US | San Diego | Accumetrics Inc | 20120627 | 1424553174.8365 | Voluntary: Firm Initiated | Class II | Not available | 003a5ea3fd3c537ecf1686e21ce6df44b366d1d8329815a9a92fcd45f86b6bdc | Letter
Z-0008-2014 | During final testing of the Draeger Fabius GS Premium, Fabius OS, and Fabius Tiro anesthesia machines, some units did not pass the high voltage test and the minimum clearance between an electrical component and the unit housing was not maintained for some from a specific batch. | Ongoing | Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Azerbaijan, Bangladesh, Brazil, Bulgaria, Canada, Chile, China, Columbia, Ecuador, Egypt, France, Germany, Greece, Hungary, India, Indonesia, Israel, Italy, Ivory Coast, Japan, Kazakhstan, Latvia, Libya, Lithuania Malaysia, Morocco, Myanmar, Nepal, Philippines, Poland, Rumania, Russian Fed., Saudi Arabia, Senegal, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom, Uruguay, Vietnam, White Russia Turkmenistan. | 1127 units | 20130814 | PA | 66025 | Devices | Draeger Fabius GS Premium, Fabius OS, Fabius Tiro, Fabius Tiro D-M used to administer anesthesia and ventilation to patients during surgical procedures. | US | Telford | Draeger Medical, Inc. | 20131016 | 1424553174.8365 | Voluntary: Firm Initiated | Class I | Part Nos. 8607000 (GS Permium) 8606000 (Fabius OS and Fabius Tiro) 8608400 (Fabius Tiro D-M)ASEC-0055 with lot numbers: ASEC-0186, ASED-0006, ASED-0074, ASED-0291, ASEC-0056, ASEC-0205, ASED-0007, SED-0075 ASED-0292, ASEC-0061, ASEC-0206, ASED-0020, ASED-0079, ASED-0293, ASEC-0062, ASEC-0207, ASED-0021, ASED-0139, ASED-0294, ASEC-0070, ASEC-0208, ASED-0023, ASED-0160, ASED-0295, ASEC-0071, ASEC-0209, ASED-0024, ASED-0161, ASED-0296, ASEC-0112, ASEC-0210, ASED-0025, ASED-0162, ASEE-0021, ASEC-0124, SEC-0211, ASED-0031, ASED-0163, ASEE-0022, ASEC-0125, ASEC-0212, ASED-0032, ASED-0164, ASEE-0023, ASEC-0126, ASEC-0213, ASED-0033, ASED-0165, ASEE-0041, ASEC-0169, ASEC-0214, ASED-0034, ASED-0166, ASEE-0042, ASEC-0170, ASEC-0236, ASED-0035, ASED-0167, ASEE-0058, ASEC-0171, ASEC-0248, ASED-0036, ASED-0168, ASEE-0060, ASEC-0172, ASEC-0249, ASED-0037, ASED-0169, ASEE-0061, ASEC-0173, ASEC-0250, ASED-0038, ASED-0203, ASEE-0065, ASEC-0174, ASED-0001, ASED-0044, ASED-0253, ASEE-0066, ASEC-0175, ASED-0002, ASED-0045, ASED-0286, ASEE-0083, ASEC-0176, ASED-0003, ASED-0046, ASED-0287, ASEE-0103, ASEC-0184, ASED-0004, ASED-0047, ASED-0288, ASEE-0136, ASEC-0185, ASED-0005, ASED-0051, ASED-0290, ASEC-0054, ASEC-0194, ASED-0247, ASED-0263, ASEE-0072, ASEC-0167, ASEC-0195, ASED-0248, ASED-0264, ASEE-0089, ASEC-0178, ASEC-0196, ASED-0254, ASED-0265, ASEE-0090, ASEC-0179, ASEC-0197, ASED-0255, ASED-0266, ASEE-0131, ASEC-0180, ASEC-0227, ASED-0256, ASED-0282, ASEE-0132, ASEC-0181, ASEC-0228, ASED-0257, ASEE-0003, ASEE-0133, ASEC-0182, ASED-0040, ASED-0260, ASEE-0017, ASEE-0134, ASEC-0183, ASED-0086, ASED-0261, ASEE-0054, ASEC-0193, ASED-0245, ASED-0262, ASEE-0071, ASED-0141, ASED-0143, ASED-0145, ASED-0147, ASED-0149, ASED-0142, ASED-0144, ASED-0146, ASED-0148, ASEB-0219 | 0053f2d69b8ca400deed7c8641c3ccd511656d8713c381f376c01900ebf61e26 | Letter
Z-1901-2013 | Reports of false abnormally high glucose readings from test strips | Ongoing | Worldwide Distribution - USA Nationwide and the countries of Chile, England , Peru, Nicaragua | *Kitted into System Kits 43435 and 45507 | 20130715 | MA | 65617 | Devices | Nova Max Glucose Test Strips (10 Count Vials) Catalog Number: 43604 These strips are kitted into Nova Max Glucose Test Strip System Kits 43435 and 45507 Product Usage: The Nova Max Glucose Test Strips are for use with the Nova Max Family of Glucose Monitors (Nova Max, Nova Max Plus, Nova Max Link, BD Logic and Paradigm Link). These monitors are intended to be used for the quantitative measurement of glucose in fresh capillary whole blood. They are intended for use by people with diabetes mellitus in the home and by healthcare professionals in clinical settings as an aid to monitor the effectiveness of diabetes control | US | Waltham | Nova Biomedical Corporation | 20130821 | 1424553174.8365 | Voluntary: Firm Initiated | Class I | Lot #/Exp.Date: 1020212032 02/28/2014 1020212153 06/30/2014 1020212207 07/31/2014 1020212291 10/31/2014 1020213074 03/31/2015 1020213109 04/30/2015 1020412255 09/30/2014 1020412341 12/31/2014 1020512087 03/31/2014 1020811350 12/31/2013 102121304 02/28/2015 | 0059457d82ee5d2780160121b6e90c112bc5159b7401bb1882b01310efc6cb97 | Letter
Z-1831-2013 | Product fails to meet the accuracy at the low end of the measuring range. | Ongoing | Distribution to CA, GA, MN, NC, UT and the District of Columbia. | 36 cartridges -100 tests per cartridge | 20130430 | CA | 65084 | Devices | uTASWako DCP, Lot TG340 2013.11, REF 995-60701, Wako Pure Chemical Industries, Ltd. 1-2 Doshomachi 3-chome, Chuo-Ku Osaka 540-8605 Japan Distributed by Wako Diagnostics Wako Life Sciences, Inc. 1025 Terra Bella Ave., Suto A. Mountain View, CA 94043 USA. The product is a single reagent cartridge run on an automated system and each assay performed on a single disposable chip using microfluidic electrophoretic separation. The reagent cartridge contains various buffers and antibodies. | US | Mountain View | Wako Life Sciences, Inc. | 20130807 | 1424553174.8365 | Voluntary: Firm Initiated | Class II | TG340, Exp. 11/30/13 | 005e5022aca369c3503792d6caa8eb63c46cf3e6c6d2da642bcd88edaaae8227 | Letter
Z-1236-2013 | The pre-vacuum steam sterilization (wrapped method) temperature in the Instructions for Use incorrectly listed 121-123 C , instead of the correct 132-133 C. | Ongoing | Worldwide Distribution: USA (nationwide) and countries of: Mexico, Portugal, Latin America, India, Greece, Taiwan, Singapore, New Zealand, Phillipines, Netherlands, France, Australia, Canada, Japan, Scandinavia, Iberia, Brasil, Romania, South Africa, Chile, Korea, China. | 147,921 - all sizes | 20130327 | CA | 64682 | Devices | Stryker Irrigator Reusable Tips: 3mm X 32cm Regular Tip, Vented, Metal Finish. Suction irrigator tips aid in cleaning and clearing the operative site during laparoscopic surgery. | US | San Jose | Stryker Endoscopy | 20130515 | 1424553174.8365 | Voluntary: Firm Initiated | Class II | Manufacturer Part Number: 250-070-490 | 0067371d737cbda5b11cd7fba9608897f103dfad09d9ca1ec90eeb72c0d2fba1 | Letter
Z-1471-2013 | Seals may disengage from the cannula which may result in a component inadvertently disengaging into the patients cavity | Ongoing | Worldwide Distribution - USA (nationwide) and internationally to the following countries: Canada Austria Bahrain Benoni Centurion Chatsworth Czech Republic Denmark Egypt Finland France Germany Ireland Israel Italy Japan Kuwait Netherlands Norway Portugal QATAR Saudi Arabia Serbia Singapore South Africa Spain OMAN Sweden Switzerland UAE United Kingdom Venezuela | 2705 units | 20130506 | CT | 65093 | Devices | Versaport Bladeless Optical Fixation Cannula - 5mm Short Product Code: ONBFCA5SH The Versaport" Bladeless Optical 5mm trocars are intended for use in a variety of gynecologic, general, thoracic and urologic endoscopic procedures to create and maintain a port of entry. | US | North Haven | Covidien LP | 20130612 | 1424553174.8365 | Voluntary: Firm Initiated | Class II | Lot Numbers: N2J0315X through N3A0389X Note: The specific lot numbers listed below are not affected by this recall and are acceptable for use: N2H0045X,N2H0166X,N2H0286X, N2H0357X,N2 H051 6UX | 008f909ff1528a3aa65475dd1aa439f98e8d091f2f5a4a20639cb0c83622f4a0 | Letter
Z-0195-2013 | The recall was initiated because SenDx Medical Inc. became aware of a software issue in the ABL80 FLEX CO-OX analyzer related to the reporting of the hemoglobin fractional parameters. | Terminated | Worldwide Distribution-- US Distribution including the state of Ohio and the countries of Switzerland, Germany, France, Poland, Austria, Italy, Japan, Australia, Denmark, China, Sweden, Norway, Spain, United Kingdom, Netherlands, and Czech Republic. | 767 units | 20100106 | CA | 63368 | Devices | ABL80 FLEX CO-OX analyzer, model #393-841 (all software versions prior to 1.35). The ABL80 FLEX CO-OX is a portable, automated analyzer that measures pH, blood gases, electrolytes, glucose and oximetry in whole blood. The ABL80 FLEX CO-OX system is intended for use by trained technologists, nurses, physicians and therapists. It is intended for use in a laboratory environment, near patient or point of care setting | US | Carlsbad | Sendx Medical Inc | 20121114 | 1424553174.8365 | Voluntary: Firm Initiated | Class II | Serial Numbers: all ABL80 FLEX CO-OX analyzers with sofware versions 1.30, 1.31, 1.32, 1.33, and 1.34. | 00a1d8ac454ba0eca3b0895cc32a183ea3577e0d15e018dd53989a562d01d9d2 | Letter
Z-1849-2012 | Software Version 3.1.1 does not meet navigational accuracy. | Terminated | Nationwide Distribution including the states of CO, IN, and MN. | 3 units | 20101012 | MA | 61741 | Devices | Medtronic 0-arm Imaging System with Software Version 3.1.1, Catalog Numbers: BI-700-00027-120R and BI-700-00027-120. The 0-arm Imaging System is a mobile x-ray system designed for 2D fluoroscopic and 3D imaging. | US | Littleton | Medtronic Navigation, Inc. | 20120627 | 1424553174.8365 | Voluntary: Firm Initiated | Class II | Serial Numbers: 125, 156, 234R. | 00aa22d018fb0c69819a21c3709379a6c7da5252913a789b724a15055a2d6050 | Two or more of the following: Email, Fax, Letter, Press Release, Telephone, Visit
Z-2011-2012 | On 05/23/2012, Stingray Surgical Products, LLC., Boca Raton, FL initiated a recall initiated due to "Instructions For Use" of Bipolar Forceps not being cleared by FDA and incorrect "Instructions For Use" . The sterilization instructions have been altered without the clearance of the FDA. | Ongoing | Worldwide Distribution - USA (nationwide) | 6326 | 20120523 | FL | 62024 | Devices | Bipolar Forceps model number: S02-XXX, S03-XXX, P02-XXX, and P03-XXX. "Instructions For Use" (QF7.5.1A.R01 thru QF7.5.1A.R06 Care and Handling of Bipolar Forceps. Bipolar Forceps designed to grasp, manipulate and coagulate selected tissue for use in general surgical procedures powered by a high frequency generator. | US | Boca Raton | Stingray Surgical Products, LLC | 20120725 | 1424553174.8365 | Voluntary: Firm Initiated | Class II | Bipolar Forceps model number: S02-XXX, S03-XXX, P02-XXX, and P03-XXX. "Instructions For Use" (QF7.5.1A.R01 thru QF7.5.1A.R06 Care and Handling of Bipolar Forceps. | 00d5b1204209b580e1173114173b20695236a2ebafa2523267ca21ea368235b5 | Two or more of the following: Email, Fax, Letter, Press Release, Telephone, Visit
Z-0165-2014 | Alcohol Prep Pads contained within the kit are not properly labeled. | Ongoing | Worldwide Distribution - USA Nationwide and the country of Israel. | 1,719 starter kits | 20130927 | MA | 66425 | Devices | OmniPod Insulin Management System Starter Kit; Model Reference Number: SKT-UST400 (Starter Kit) Product Usage: The Starter Kit (SKT-UST400) is provided to users as an introduction to the OmniPod Insulin Management System prior to initiating care. It includes the Personal Diabetes Manager (PDM) with two AAA alkaline batteries, a USB cable, a User s Guide, a Soft Carrying Case, a Troubleshooting guide, a Warranty Registration Card, and two Alcohol Pads. The OmniPod Insulin Management System is intended for subcutaneous delivery of insulin at set and variable rates for the management of diabetes mellitus in persons requiring insulin and for the quantitative measurement of glucose in fresh whole capillary blood (in vitro) from the finger. The glucose measurements should not be used for the diagnosis of or screening for diabetes. The PDM glucose meter is intended for single-patient use and should not be shared. Abbott FreeStyle test strips are used with the built-in FreeStyle meter for the quantitative measurement of blood glucose in fresh whole capillary blood from the finger, upper arm and palm. Abbott Freestyle Control Solutions are used to verify that the meter and test-strips are working together properly and that the test is performed correctly. | US | Bedford | Insulet Corporation | 20131113 | 1424553174.8365 | Voluntary: Firm Initiated | Class II | Lot Numbers: L13483, L13485, L13478, & L13481 | 00e29a64451148fb317b8a91e05ca60bc28e541ab43334fddfabbabafc1044b8 | Letter
Z-1309-2013 | Possible damage to the sterile packaging (blister tray) of some Karl Storz sterile products, which may result in airborne contamination. | Terminated | Distributed Nationwide and in Canada. | 9432 units | 20090601 | CA | 55358 | Devices | Electrode, Roller Ball, 24FR (27050NK/6). Resectoscope accessories (ie. electrode, cutting loops and cold knives). | US | El Segundo | Karl Storz Endoscopy America Inc | 20130522 | 1424553174.8365 | Voluntary: Firm Initiated | Class II | Lot 36833 | 00e6e0cccb3024c6c1a98b7bbcd3fe2eef31d06fe3d19c8b62cecc3a07734293 | Letter
Z-0313-2013 | The product may mislabel patient information on culture specimen bottles due to a software miscommunication. | Ongoing | Worldwide Distribution - USA including AR, AZ, CA, FL, KS, LA, MD, MN, MT, NC, NV, NY, OR, PA, RI, SC, TX, VA and WA. Internationally to Argentina, Armenia, Austria, Bangladesh, Bosnia and Herzegowina, Brazil, Bulgaria, Burkina FASO, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, COTE DIVOIRE, CROATIA, CZECH REPUBLIC, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, FRANCE, FRENCH POLYNESIA, GERMANY, GREECE, GUATEMALA, HONDURAS, HONG KONG, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LATVIA, LITHUANIA, MALTA, MEXICO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, PANAMA, PERU, POLAND, PORTUGAL, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, UNITED STATES, URUGUAY, VENEZUELA, VIETNAM, PALESTINE, AND SERBIA. | 250 kits | 20121016 | MO | 63467 | Devices | bioMerieux Kit, USB HDW, BTA3D CONTROL-B-B.40, Clinical instrument software, USB update and mounting hardware-English, catalog number 411958 Software update for the clinical instrument which is most commonly used to isolate and recover specimens drawn from adult and older pediatric patients that are acutely ill with signs and symptoms suggestive of blood stream infection, which can be due to a variety of sources including lung, abdomen, urinary tract, endovascular, and skin. | US | Hazelwood | Biomerieux Inc | 20121128 | 1424553174.8365 | Voluntary: Firm Initiated | Class II | Identified by B.40 firmware | 010b57eb148fdd91d55eea2f048700cf3c42f4ead2e552246c8f63156dbea31c | E-Mail
Z-0452-2014 | Three lots of Titanium Trochanteric Fixation Nail (TFN) System are missing the second bend in the nail. | Ongoing | Nationwide Distribution including IL, CA, AZ, FL, TX, UT, VA, NY, NM, and MT. | 17 | 20130730 | PA | 66417 | Devices | Titanium Trochanteric Fixation Nail (TFN)System Intended to treat stable and unstable fractures of the proximal femur. | US | West Chester | Synthes USA HQ, Inc. | 20131211 | 1424553174.8365 | Voluntary: Firm Initiated | Class II | Part # 456.391S with lot numbers 5536930, exp. 5/20/2016, and 7161401, exp. 12/31/2021. | 010f5013abf7a5fb1590cbc2c7229b0f766b6fecebeea1ae3a8e8cea36e8f740 | Letter
Z-0612-2013 | It has been determined that, in some cases, pulley drive components have broken parts or have become loose, mispositioned, or disconnected from their drive shafts. When these defects occur, the pulley may not turn at all, or may turn at an incorrect speed, resulting in inadequate washing or mixing of the contents of the reaction vessel. | Ongoing | Worldwide Distribution-USA (nationwide) and the country of Canada. | 6,858 units | 20121101 | CA | 63632 | Devices | Access Immunoassay System, Access 2 Immunoassay System, Synchron LXi 725 Clinical System, UniCel DxC 600i Synchron Access Clinical System Part Numbers: Access 81600; Access 2 81600N, 386220, A69186, A25656, A12900; Access 2 (LXi) A15642; Access 2 Section, DxC 600i A25640. The Access Immunoassay System is an in vitro diagnostic device used for the quantitative, semi-quantitative or qualitative determination of various analyte concentrations found in human body fluids. The Synchron LXi and the UniCel DxC 600i are Access systems integrated with a chemistry analyzer. | US | Brea | Beckman Coulter Inc. | 20130102 | 1424553174.8365 | Voluntary: Firm Initiated | Class II | Part Numbers: Access 81600; Access 2 81600N, 386220, A69186, A25656, A12900; Access 2 (LXi) A15642; Access 2 Section, DxC 600i A25640 | 01128cd5aed0b85a0821a4ba68a95226b73e87eb2736d8b6cc9f7d58eeb7c5d5 | Letter
Z-2050-2013 | These products are being recalled because they were distributed without adequate labeling and instructions for use. Inadequate labeling could lead to misuse of the Test Kits and incorrect test results. | Ongoing | Worldwide Distribution - US Nationwide in the country Virgin Islands | various quantities | 20130729 | AZ | 65841 | Devices | SARKEN STAT! 1- Panel Test Kit-THC Product Usage- Labeled as FOR FORENSIC USE ONLY (Employee, Student or Probation Testing) . | US | Tempe | Sarken, Inc. | 20130904 | 1424553174.8365 | Voluntary: Firm Initiated | Class II | Manufacturer's Product Number DTH-114 Lot Number THC1120052 THC2120005 THC211100044 THC2100044 THC2120006 Expiration Date: 12/2013 10/2014 09/2014 09/2014 11/2014 | 0116131a6d33b8a6f6fdd12b6c66912d2215d66e12abda0f8da82e30965d524c | Letter
Z-1812-2013 | In vitro diagnostic test kit may exhibit invalid internal negative control results. | Ongoing | Nationwide Distribution | 177 kits | 20130613 | MD | 65619 | Devices | BD Affirm VPIII Microbial Identification Tests, packaged in kits, 120 test kit, Catalog # 446257 and 24 test kit, Catalog # 446252; kits are labeled in part ***Becton, Dickinson and Company, 7 Loveton Circle Sparks, MD 21152 USA, 800-638-8663, www.bd.com/ds*** The Affirm VPIII Microbial Identification Test is a DNA probe test intended for use in the detection and identification of Candida species, | US | Sparks | Becton Dickinson & Co. | 20130807 | 1424553174.8365 | Voluntary: Firm Initiated | Class II | Cat #446257: Lot 3035126 Exp 12/04/13 Cat #446252: Lot 3065261 Exp 08/31/13 | 0116753f69d91c96ba3de2ccb326115684c90e2f27385765d3c4d0a372374430 | Letter
Z-1899-2012 | The Y adapter within the breathing circuit may spontaneously crack, causing a leak in the closed ventilation system. This leak could lead to a decrease in the tidal volume delivered to the ventilated patient. | Ongoing | Worldwide distribution: USA (nationwide) including Puerto Rico and internationally to: Canada, Egypt, France, Germany, Great Britain, Italy, Kuwait, Mexico, Portugal, Saudi Arabia, Spain, Taiwan and United Arab Emirates. | null | 20120529 | IL | 61943 | Devices | AirLife Isothermal Breathing Circuit, Infant Respiratory Circuit, Heated; an Rx device for single patient use; CareFusion, 22745 Savi Ranch Parkway, Yorba Linda, CA 92887 USA, Made in Mexico Respiratory breathing circuits are used with a ventilator when mechanical ventilation is administered to a patient. | US | McGaw Park | Carefusion 2200 Inc | 20120711 | 1424553174.8365 | Voluntary: Firm Initiated | Class I | product codes 10127-4H1, 10331N-4S2, 10351-4H2, 10387-4S2, 10555-4S2, 10706-4S2, 10790-4S2, 10814-4S2, 10849-4S2, 1175-4H2, 1208-4H1, 1553-4H1, 1815-409, 1998-4H1, 2119-4S2G, 2120-4S2G, 4319-4H2, 4326-4H2, 4652-5014677-4H2, 4875-4H2, 5665-4H2, 5682-409, 5732-4H1, 5868-4H2, 6002-4H2, 6040-501, 6074-4H1, 6198-4H2, 6274-H12, 6301-H12, 6313-501, 6459-4H2; 6485-4H2, 6588-4H2, 6942-4H2, 7100-4S2, 7147-4S2, 7156-4S2, 7184-4S2, 7206-4S2, 7213-4S2, 7354-4S2, 7395-4S2, 7407-4S2, 7441-4S2, 7483-4S2, 7526-4S2, 7575-4S2, 7582-4S2, 7588-4S2, 9053-4S2, 9054-4S2, 9055-4S2, 9056-4S2, 9058-4S2, RC41-12027, RC41-18039, RC41-18052, RC41-18060, RC51-12074, RC51-18004, RT4851-00, RT4851-12 and RT4851-18; includes all lot numbers of the affected product codes manufactured prior to February 3, 2012, lot #0000414257 and any lot numbers beginning with the letter Y. | 012487ace80ef6a3d6de424003d476b7f47867cc8cbb4866b38ea411b6c5f508 | Letter
Z-0754-2013 | It was reported by a customer site that when using GE AdvantageSim MD on Advantage Workstation, the series could be incorrectly labeled in image view when multiple series of an exam are loaded simultaneously in Advantage Sim MD and if their series dates are different. A mismatch of series label for structure sets may lead to under-treatment of a tumor due to too small coverage of volume treated. A second issue was discovered internally in which the Interpolation in Advantage Sim MD 7.5 would not give a correct result when some part wee removed from an existing structure. If the contour truncation is not recognized, it may lead to inappropriate irradiation of the patient. | Ongoing | Worldwide Distribution - USA including AL, CT, FL, IL, MA, MO, NJ, NY, PA, TX VA, WA, WI and Internationally to TURKEY, TAIWAN, SWITZERLAND, SWEDEN, SOUTH AFRICA, NIGERIA, MALAYSIA, KOREA, JAPAN, INDIA, GERMANY, FRANCE, EGYPT, DENMARK, CHINA, CHILE, CANDA, BRAZIL, AUSTRIA, and GREAT BRITAIN. | 65 | 20120524 | WI | 61611 | Devices | AdvantageSim MD versions 7.4 through 7.6, Model 5160092-2 Version 7.4, 5160092-3 Version 7.5, 5160092-4 version 7.6. | US | Waukesha | GE Healthcare, LLC | 20130206 | 1424553174.8365 | Voluntary: Firm Initiated | Class II | Mfg Lot or Serial # 00000L022E05AE 00000139595HP3 000000010003GS 000000410012GS 000000210016GS 000000210013GS 00000138998HP0 00000139435HP2 000000410010GS 000000410009GS 000000210005GS 000000210011GS 000000210012GS 000000210015GS 00000138991HP5 000000210014GS 000000310005GS 00000139574HP8 00000139604HP3 00000L026CC222 000000210010GS 000000010004GS 000000410011GS 000000410007GS 000000210023GS 000000210025GS 000000210026GS 000000210022GS 000000210024GS 000000410005GS 000000410006GS 000000310003GS 000000310002GS 000000210006GS 000000210007GS 000000210003GS 000000410004GS 000000310001GS 000000010001GS 000000210030GS 000000210008GS 000000410016GS 000000P8059001 000000210027GS 000000410008GS 000000410001GS 000000210017GS 000000210019GS 000000210021GS 000000210020GS 000000310008GS 000000210018GS 000000410014GS 000000210029GS 000000210004GS 000000210001GS 000000210009GS 000000410003GS 000000410013GS 000000210028GS 000000410017GS 000000410015GS 000000310007GS 000000310006GS 000000010005GS | 014e3f3e228b35629db3c5ecb6aa8055bad45b5d6000cc9b7c2f593fb92cd2c8 | Letter
Z-1766-2013 | Blood blushing/leaking from the surface of the graft after implantation | Ongoing | Worldwide distribution: USA including PA and countries of: BRAZIL, BULGARIA, FRANCE, GEORGIA, GERMANY, GREECE, ITALY, LEBANON, NETHERLAND, POLAND, SPAIN, and TURKEY. | 2 units | 20130617 | MA | 65553 | Devices | AlboGraft Woven with Collagen, Straight 15x30 Catalog Number: ATC1530 AlboGraft Vascular Grafts are made of synthetic material and designed to replace sections of damaged or malfunctioning arteries. | US | Burlington | LeMaitre Vascular, Inc. | 20130731 | 1424553174.8365 | Voluntary: Firm Initiated | Class I | Lot number/Expiration Date 56890A 2015-04 | 01521e1a49cbd2ec2070d36292b2521f00962d6dda6014139a9eb90db44b4c7a | Two or more of the following: Email, Fax, Letter, Press Release, Telephone, Visit
Z-1675-2013 | The RITA Main Cable product contains the incorrect Instructions for Use. | Ongoing | Worldwide Distribution, including the US, Austria, China, Italy, Russia and the UK. | 13 units | 20130611 | GA | 65439 | Devices | AngioDynamics RITA Main Cable, 9 Pin to 14 Pin, REF 700-101892. RITA Main Cables are indicated for use with the RITA RF Generators and RITA electrosurgical devices. | US | Manchester | Angiodynamics, INC | 20130717 | 1424553174.8365 | Voluntary: Firm Initiated | Class II | Catalog #700-101892, Lot number: 30.12 | 0180bba41bb502a50b6df06a22e3e80c3e64d5cd66c4abe58b7999d67d4c318c | Letter
Z-2254-2012 | Hitachi discovered a software error with a feature called Multi-Planar Reconstruction (MPR). MPR takes an original volume dataset reconstructed in any plane view and creates a new image set from another view (i.e. axial to coronal, coronal to sagittal, etc.). The error causes the first image in the dataset to have a right/left reversed orientation. The software error can occur randomly when executing the MPR task. | Ongoing | US Nationwide Distribution - including the states of:: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NJ, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WI, WY and Mexico. | 165 units | 20120615 | OH | 62870 | Devices | Hitachi Echelon MRI System/Oasis MRI, Software. These are software controlled MRI devices. Product Usage: The Echelon and Oasis MR systems are imaging devices, intended to provide the physician with physiological and clinical information that can be useful in diagnosis determination. | US | Twinsburg | Hitachi Medical Systems America Inc | 20120829 | 1424553174.8365 | Voluntary: Firm Initiated | Class II | Echelon, V001 through V061, inclusive. Oasis, M001 through M105, inclusive, M951 | 019e11a1fafe6dd64c707f23cb09f467081201541ea1ac54fcc8cbec63523c88 | Letter
Z-0199-2014 | Anspach Effort, Inc. Palm Beach Gardens, FL initiated a correction of its medical Device labeling for its Pneumatic Systems Operating Manual, Cutting Burr Chart and Product Catalog due to a recent review of their regulations and compliance requirements that determined the claim that the device is capable of dissection of biometals, bioplastics and cement has been verified, but is is not part of the currently cleared indications for use fo the device(s). The currently cleared indications for use for the pnuematic and electric motor systems is "cutting and shaping bone". | Ongoing | Worldwide Distribution: USA Nationwide and the countries of: Egypt, India, Germany, Turkey, Canada, Saudi Arabia, Greece, United Kingdom, Cayman Islands, Kazakhstan, Republic of Korea, Taiwan, Hong Kong, China, Lebanon, Argentina, Denmark, Belgium, Netherlands, Mexico, France, Australia, Panama, Italy, Guam, United Arab Emirates, Israel, Polan, Venezuela, Austria, Hungary, Lithuania, Norway, Malaysia, Colombia, Costa Rica, Ecuador, Jamaica, Peru, Thailand, Italy, Czech Republic, Cameroon, South Africa, Portugal, Finland, Sweden, Nigeria, Bahamas, Cambodia, Dominican Republic, Switzerland, Chile, Singapore, Spain, Syrian Arab Republic, Jordan, Japan. | 2020 | 20110506 | FL | 65863 | Devices | Cutting Bur Chart (part number 20-0019) utilized with the Electric and Pneumatic Motor Systems including: MicroMax Motor Systems, XMax Motor Systems, EMax2 Motor Systems, and Emax2Plus Motor Systems, all models and serial numbers. Product Usage: Indication for use is cutting and shaping bone | US | Palm Beach Gardens | The Anspach Effort, Inc. | 20131120 | 1424553174.8365 | Voluntary: Firm Initiated | Class II | part number 20-0019 | 01ab0fc3517d135d3dd6d28d9a7d616796a48e5070f662075cb244616ba9efe5 | Letter